## Coordinate Meas Machine
|  |  |  | Doc No : | ISW-QA-032
|  |  |  | Issue No : | 4
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.07.10
Coordinate Measuring Machine (Crysta 574)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of all Complex Three Dimensional Measurements for all Critical Inspections Loactions
which is either not possible or very difficult to check with normal Routined Measuring Equipments
Reference Standards/Materials : Master Ball (19.9921 mm)
Description: | Machine makes use of Geopack Software for basic evaluation & measurements of various
geometrical parameters. Machines could be programmed for individual components Inspection
Locations with the availability of Base Drawing as a reference.It is possible to check various
geometrical dimensions and analysis of the same can be done as desired & needed for Inspection.
Machine has a facility for both touch trigger probing and Scanning .
Standard Accessories : |  | The standard accessries required for the operation of the machine are Ruby Stylus,
Magnetic V Block,Master Ball (19.9921 mm)
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.In case
of any doubt in the overall performance of the machine,the periodic calibration is done
using a Straight Test Bar.The Test Bar is subjected to Repeat checking of Diameter 25 Nos
of Times and checked for repeatability.
Calibration Frequecy: In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents : |  | Operation Standard "Annexure to ISW-QA-032"
Prepared By |  | Checked By | Approved By |  | Released By

## Coordinate Measuring Machine
[6 empty] | OPERATION STANDARD | [5 empty] | CONTROL NUMBER
SOP-QA-041
Equipment Name | Coordinate Measuring Machine (Crysta 776) |  |  | Process | Geometrical Dimensions Checking |  | Section | Date | 2014-06-01 |  | PREPARED |  | CHECKED | APPROVED
Components | Shaft,Gears & Cam shafts |  |  | Oprn. No. |  |  | Quality Assurance | Issue No | 01
[8 empty] | Rev.NO | 00
Picture |  |  | S. No. | Operation Procedure |  | Vital Point Of Operation |  | Things which You should Not Do |  |  | Result in case you do it
|  |  | 1 | Switch On Power Supply To Instrument |  | Perform With Dry Hands |  | Don't operate With Wet Hands |  |  | May Get Electric Shock
Computer & Geopack Software Options
|  |  | 3 | Visually Check Condition Of Renisaw Probe |  | Condition Should Be OK |  | If NotOK Report Shift Incharge |  |  | Error In Measurment Will Occur
|  |  |  | Head & Ruby Stylus |  |  |  | Don't Operate With Damaged Stylus
|  |  | 4 | Bring the Machine to Home Position by
bringing the Detector Unit towards extreme
End of Slides as shown in the fig
|  |  | 5 | For Probe Calibration, Select "Create a New
Part" Graphical Interface Button
|  |  | 6 | Click on "CMM Learn Mode" Graphical
Interface Button
|  |  | 7 | Enter Expansion Coefficient [11.7] Armco Iron
|  |  | 8 | Click on "Probe Builder" & Select PH10M,
SP600M & Reqd Dia Stylus Graphical Interface
Main Machine |  |  |  | Button
|  |  | 9 | Click on "Calibrate" Button and proceed as
per the message on the Calibration Screen.
Probe Hit on the component or Master Probe
is operated by Hand Operated Toggle Switch
|  |  | 10 | After Probe getting Calibrated, Select the |  | Program called should
|  |  |  | Program based on the component ,model |  | be specific to the checking
Home Position | Calibration Screen |  |  | and stage viz Hobbing,Shaving or Final Stage |  | Component,model & Stage
as appearing in the Part Selection Screen
|  |  | 11 | Position the Component on the Magnetic
V Block as shown in the photogrpahs
displayed against the component,model
& Stage Selected
|  |  | 12 | Check for the right Stylus Diameter and then |  | Check For Right Stylus |  | Always Use Specified Stylus |  |  | Stylus & Probe Head May Get Damaged
define plane ,origin & Axis as per the stated
message & photographs displayed against
the component,model & Stage Selected
Renisaw Probe Head & Stylus
|  |  | 13 | In case there is many components to be
checked ,Call "Repeat Measurements "
Graphical Interface Button once the Program
Cycle is completed
|  |  | 14 | In Case of any abnormal selection of Program
Mistake/ Collision Foreseen,Press "Emergency
Button" Identified as Red in Color on the Hand
Operated Toggle Switch to stop the Machine
Hand Operated Toggle Switch |  |  |  | Operation
SAFETY DEVICES TO BE USED |  |  | ATTENTION |  |  |  |  | WORKING HEADS |  |  |  | REMARKS
Always Change STYLUS Carefully. Refer To Manual Provided With Instrument (In Case Of Any Problem)
Name of safety device | Qty | Remarks | Route to be followed in case of any trouble |  |  |  |  | 1. MACHINE STARTUP |  | 5. MACHINE STOP
|  |  | A Shift, General Shift |  | Line Asso. Exec. Sr. Exe. A.M D.M. |  |  | 2. MACHINE OPERATION |  | 6 CLEANING
3. QUALITY MEASUREMENT
|  |  | B & C Shift |  | Line Asso. Exec. Sr. Exe.
VITAL CHECK POINTS ARE SHOWN IN RED COLOUR
4. STYLUS CHANGE

## cmm 8.5.2017
[6 empty] | OPERATION STANDARD | [5 empty] | CONTROL NUMBER
SOP-QA-041
Equipment Name | Coordinate Measuring Machine (Crysta 776) |  |  | Process | Geometrical Dimensions Checking |  | Section | Date | 8.5.2017 |  | PREPARED |  | CHECKED | APPROVED
Components | Shaft,Gears & Cam shafts |  |  | Oprn. No. |  |  | Quality Assurance | Issue No | 2 |  | shivananda |  | prashanth | sathoshi san
[8 empty] | Rev.NO | 00
Picture |  |  | S. No. | Operation Procedure |  | Vital Point Of Operation |  | Things which You should Not Do |  |  | Result in case you do it
|  |  | 1 | Switch On Power Supply To Instrument |  | Perform With Dry Hands |  | Don't operate With Wet Hands |  |  | May Get Electric Shock
Computer & Geopack Software Options
|  |  | 3 | Visually Check Condition Of Renisaw Probe |  | Condition Should Be OK |  | If NotOK Report Shift Incharge |  |  | Error In Measurment Will Occur
|  |  |  | Head & Ruby Stylus |  |  |  | Don't Operate With Damaged Stylus
|  |  | 4 | Bring the Machine to Home Position by
bringing the Detector Unit towards extreme
End of Slides as shown in the fig
|  |  | 5 | For Probe Calibration, Select "Create a New
Part" Graphical Interface Button
|  |  | 6 | Click on "CMM Learn Mode" Graphical
Interface Button
|  |  | 7 | Enter Expansion Coefficient [11.7] Armco Iron
|  |  | 8 | Click on "Probe Builder" & Select PH10M,
SP600M & Reqd Dia Stylus Graphical Interface
Main Machine |  |  |  | Button
|  |  | 9 | Click on "Calibrate" Button and proceed as
per the message on the Calibration Screen.
Probe Hit on the component or Master Probe
is operated by Hand Operated Toggle Switch
|  |  | 10 | After Probe getting Calibrated, Select the |  | Program called should
|  |  |  | Program based on the component ,model |  | be specific to the checking
Home Position | Calibration Screen |  |  | and stage viz Hobbing,Shaving or Final Stage |  | Component,model & Stage
as appearing in the Part Selection Screen
|  |  | 11 | Position the Component on the Magnetic
V Block as shown in the photogrpahs
displayed against the component,model
& Stage Selected
|  |  | 12 | Check for the right Stylus Diameter and then |  | Check For Right Stylus |  | Always Use Specified Stylus |  |  | Stylus & Probe Head May Get Damaged
define plane ,origin & Axis as per the stated
message & photographs displayed against
the component,model & Stage Selected
Renisaw Probe Head & Stylus
|  |  | 13 | In case there is many components to be
checked ,Call "Repeat Measurements "
Graphical Interface Button once the Program
Cycle is completed
|  |  | 14 | In Case of any abnormal selection of Program
Mistake/ Collision Foreseen,Press "Emergency
Button" Identified as Red in Color on the Hand
Operated Toggle Switch to stop the Machine
Hand Operated Toggle Switch |  |  |  | Operation
SAFETY DEVICES TO BE USED |  |  | ATTENTION |  |  |  |  | WORKING HEADS |  |  |  | REMARKS
Always Change STYLUS Carefully. Refer To Manual Provided With Instrument (In Case Of Any Problem)
Name of safety device | Qty | Remarks | Route to be followed in case of any trouble |  |  |  |  | 1. MACHINE STARTUP |  | 5. MACHINE STOP
|  |  | A Shift, General Shift |  | Line Asso. Exec. Sr. Exe. A.M D.M. |  |  | 2. MACHINE OPERATION |  | 6 CLEANING
3. QUALITY MEASUREMENT
|  |  | B & C Shift |  | Line Asso. Exec. Sr. Exe.
VITAL CHECK POINTS ARE SHOWN IN RED COLOUR
4. STYLUS CHANGE