## Cam Tester
[5 empty] | Doc No : | ISW-QA-030
[5 empty] | Issue No : | 4
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 01.07.10
Cam Tester
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of Cam Profile of Various Cam Components..
Reference Standards/Materials : |  |  | Master Cam (MF1001-3-F008),Height Master Gauge(MF1001-4-G007),
Gauge Block,Electronic Dial Unit
Description: | This Machine could be used effectively for checking of various cam profile in both manual & automatic
mode.It is possible to program the machine for a specific component for different values of lift angle
and lift error as specified and could be saved with a distinguished identification No.The whole software
program is menu driven thereby making the machine convenient for use.The detailed descriptions of
the various menus can be referred from Operating Manual.
Standard Accessories: |  | The standard accessries required for the operation of the machine are Height Master Gauge,
Gauge Block , Electronic Dial Unit.
Calibration: | The annual calibration of the instrument is subject to satisfactory performance of the machine.
The master Cam piece are checked for Cam Profile on Periodic Basis and the results are compared
with the Master Graph taken during the Machine Installation time.
Calibration Frequecy: In case of suspected performance ,the instrument is calibrated by the maker .
Referred Document : Operation Standard "Annexure to ISW-QA-030"
Prepared By |  | Checked By |  | Approved By |  | Released By

## CamProfileTester
OPERATION STANDARD
Annexure To ISW-QA-030
Equipment Name |  | Cam Profile Tester |  |  | Process | Cam Profile Mesurment |  | Section | Date | 21st Dec 05 |  |  | PREPARED |  | CHECKED | APPROVED
Components |  | All Cam Components |  |  | Oprn. No. |  |  | Quality Assurance | Iss.no. | 3 | Rev. | 0
|  |  |  | S. No. | Operation Procedure |  | Vital Point Of Operation |  | Things which You should Not Do |  |  |  | Result in case you do it
|  |  |  | 1 | Switch On UPS
|  |  |  | 2 | Switch On Power Supply To Printer, Plotter, Personal Computer & Cam Tester M/c Sequentially |  | Check For Power Supply |  | Don't Run Computer On UPS |  |  |  | UPS May Get Trip
|  |  |  | 3 | Push Master\Motor Button & Select Manual Mode
|  |  |  | 4 | Now Fix The suitable Work Driver & Driven Jig To Component & M/c Center Depending on Job. |  | Maintain Distance Between Stylus & Driving Jig |  | Stylus Should Not Hit By Job In Any Case |  |  |  | Stylus May Get Damage
Cam Tester Software Icon |  | List of comp details feed in the software
Slightly Tight Job In Driven Jig
|  |  |  | 5 | Clamp Job In Driven Jig. Locate Driven Jig At M/c Center Locking With Extended Bar Of Driving Jig & Tight Slightly With Hexagonal Key Provided |  | In Case Of Shaft Comp Cam Lock Pin Must Be Opposite To Locking Sceews Of Driven Jig
Maintain optimum distance between driven jig and Stylus to avoid any damage to follower/stylus.
Screen showing selected component details |  | Control Panel
|  |  |  | 6 | Unlock Sliding Screw, Move Stylus Close To Job & Lock It |  | Maintain Distance Between Job & Stylus
|  |  |  | 7 | Double Click On 'Cam Tester' Icon, Click & Select Work No. & Name as the case may be i.e. Pin Origin-KSP, No.-1 Cam OriginKSP,Pin Origin KRB |  | Work No. & Name Selection To Be Done Carefully
Knock Pin & Tightning Screw Opposite To Each Other |  | Electronic Display Unit |  | 8 | Click 'MAKE NOMINAL' , Check 'Cam Position' & Then Press'End' |  | Check Carriage\Stylus Position by moving follwer carraige manually to the cam piece and check the feeded value of carraige position for both cam piece In and Ex are same.
|  |  |  | 9 | Now Push Carriage Position Button Left/Right To Verify Cam Poition
|  |  |  | 10 | In Case Of Pin Origin Place Slip(Of Thickness Equal To Radius Of LockPin) On Height Master Gauge & With Help Of Electronic Dial Unit Make High Point Of Pin & Height Master Gauge At Same Lavel
|  |  |  | 11 | In Case Of No-1 Cam Origin, Make Lift & Angle Zero At Highest Point Of EX Cam Piece
Cam Tester M/c Setup
|  |  |  | 12 | Now Go To 'Measurment Mode', Click 'Measure And Report' Then 'END' & Finally Click 'START' Serially, to Initiate Measurment
|  |  |  | 13 | Push 'Auto' Button, Then 'Measure' & Carriage Will Move To Cam After Reaching Home Position
|  |  |  | 14 | Put Paper On Plotter & Press Chart Hold Button
|  |  |  | 15 | After Finishing Measurment Carriage Will Come To Home Position . Printer & Plotter Will Start Printing |  | During Home Position Setup Probe Should not Hit Job
|  |  |  | 16 | Push Manual Button, Using Sliding Screw Move Probe Away From Job, Loose Locking Srews & Unclamp The Job
17
SAFETY DEVICES TO BE USED |  |  |  | ATTENTION |  |  |  |  | WORKING HEADS
Name of safety device |  | Qty | Remarks | Route to be followed in case of any trouble |  |  |  |  | 1. MACHINE STARTUP |  |  | 5. MODEL CHANGE
[9 empty] | 2. MACHINE OPERATION |  |  | 6. MACHINE STOP
|  |  |  | A Shift, General Shift |  | Line Asso. Exec. Sr. Exe. A.M D.M.
[9 empty] | 3. QUALITY MEASUREMENT |  |  | 7. CLEANING
|  |  |  | B & C Shift |  | Line Asso. Exec. Sr. Exe. | [7 empty] | VITAL CHECK POINTS ARE SHOWN IN RED COLOUR
4. TOOL CHANGE

## CamProfileTester26.07.2015
OPERATION STANDARD
SOP-QA-40
Equipment Name |  | CAM TESTER |  |  | Process | CAM PROFILE & ANGLE MEASUREMENT |  | Section | Date | 2014-06-01 |  | PREPARED |  | CHECKED | APPROVED
Components |  | All Cam Components |  |  | Oprn. No. |  |  | Quality Assurance | Issue No | 01
[9 empty] | Rev.NO | 00
|  |  |  | S. No. | Operation Procedure |  | Vital Point Of Operation |  | Things which You should Not Do |  |  | Result in case you do it
MAIN MACHINE
|  |  |  | 1 | Switch On UPS
|  |  |  | 2 | Switch On Main Power Supply , CPU & MASTER ON Cam Tester M/c Sequentially |  | Check For Power Supply in local opearation box power lamp |  | Don't Run Computer On UPS |  |  | UPS May Get Trip
|  |  |  | 3 | Clean the Head stock center & Tail stock center by using fine cloth |  | check condition of the centers
|  |  |  | 4 | Now Fix The suitable Work Driver & Driven Jig To Component & M/c Center Depending on Job. |  | Maintain Distance Between Stylus & Driving Jig |  | Stylus Should Not Hit By Job In Any Case |  |  | Stylus May Get Damage
Slightly Tight Job In Driven Jig
|  |  |  | 5 | Clamp Job In Driven Jig. Locate Driven Jig At M/c Center Locking With Extended Bar Of Driving Jig & Tight Slightly With Hexagonal Key Provided |  | In Case Of Shaft Comp Cam Lock Pin Must Be Opposite To Locking Sceews Of Driven Jig
Maintain optimum distance between driven jig and Stylus to avoid any damage to follower/stylus.
|  |  |  | 6 | Unlock Sliding Screw, Move Stylus Close To Job & Lock It |  | Maintain Distance Between Job & Stylus
|  |  |  | 7 | Double Click On 'Cam Tester' Icon, Click & Select Work No. & Name as the case may be i.e. Pin Origin-KSP, No.-1 Cam OriginKSP,Pin Origin KRB |  | Work No. & Name Selection To Be Done Carefully
Cam tester software icons |  | program details feed in software |  | 8 | Click 'MAKE NOMINAL' , Check 'Cam Position' & Then Press'End' |  | Check Carriage\Stylus Position by moving follwer carraige manually to the cam piece and check the feeded value of carraige position for both cam piece In and Ex are same.
do not put your hand into the moving parts during operation
|  |  |  | 9 | Now Push Carriage Position Button Left/Right To Verify Cam Position
|  |  |  | 10 | In Case Of Pin Origin Place Slip(Of Thickness Equal To Radius Of LockPin) On Height Master Gauge & With Help Of Electronic Dial Unit Make High Point Of Pin & Height Master Gauge At Same Level
Measuring report screen |  | Adjustment screen
|  |  |  | 11 | In Case Of No-1 Cam Origin, Make Lift & Angle Zero At Highest Point Of EX Cam Piece
|  |  |  | 12 | Now Go To 'Measurment Mode', Click 'Measure And Report' Then 'END' & Finally Click 'START' Serially, to Initiate Measurment
|  |  |  | 13 | Push 'Auto' Button, Then 'Measure' & Carriage Will Move To Cam After Reaching Home Position
Knock Pin & Tightning Screw Opposite To Each Other |  | Electronic Display Unit |  | 14 | Put Paper On Plotter & Press Chart Hold Button
|  |  |  | 15 | After Finishing Measurment Carriage Will Come To Home Position . Printer & Plotter Will Start Printing |  | During Home Position Setup Probe Should not Hit Job
|  |  |  | 16 | Push Manual Button, Using Sliding Screw Move Probe Away From Job, Loose Locking Srews & Unclamp The Job
17
SAFETY DEVICES TO BE USED |  |  |  | ATTENTION |  |  |  |  | WORKING HEADS
Name of safety device |  | Qty | Remarks | Route to be followed in case of any trouble |  |  |  |  | 1. MACHINE STARTUP |  | 5. MODEL CHANGE
[9 empty] | 2. MACHINE OPERATION |  | 6. MACHINE STOP
|  |  |  | A Shift, General Shift |  | Line Asso. Exec. Sr. Exe. A.M D.M.
[9 empty] | 3. QUALITY MEASUREMENT |  | 7. CLEANING
|  |  |  | B & C Shift |  | Line Asso. Exec. Sr. Exe. | [6 empty] | VITAL CHECK POINTS ARE SHOWN IN RED COLOUR
4. TOOL CHANGE

## CamProfileTester26.07.2015 (2)
[7 empty] | OPERATION STANDARD | [5 empty] | CONTROL NUMBER SOP-QA-40
Equipment Name |  | Cam Tester |  |  | Process | CAM PROFILE & ANGLE MEASUREMENT |  | Section | Date | 2014-06-01 |  | PREPARED |  | CHECKED | APPROVED
Components |  | ALL Cam shaft |  |  | Oprn. No. |  |  | Quality Assurance | Issue No | 01
[9 empty] | Rev.NO | 00
|  |  |  | SI.No. | Operation Procedure |  | Vital Point Of Operation |  | Things which You should Not Do |  |  | Result in case you do it
MAIN MACHINE
|  |  |  | 1 | Switch On UPS
|  |  |  | 2 | Switch On Main Power Supply , CPU & MASTER ON Cam Tester M/c Sequentially |  | Check For Power Supply |  | Don't Run Computer On UPS |  |  | UPS May Get Trip
|  |  |  | 3 | Clean the Head stock center & Tail stock center by using fine cloth |  | check condition of the centers
|  |  |  | 4 | install the component to the work driving dog rotator |  | Maintain Distance Between Stylus & Driving Jig |  | Stylus Should Not Hit By Job In Any Case |  |  | Stylus May Get Damage
Slightly Tight Job In Driven Jig
|  |  |  | 5 | Move the tailstock to a point where it can clamp the workpiece, lock the tailstock. |  | In Case Of Shaft Comp Cam Lock Pin Must Be Opposite To Locking Sceews Of Driven Jig
Maintain optimum distance between driven jig and Stylus to avoid any damage to follower/stylus.
|  |  |  | 6 | Unlock Sliding Screw, Move Stylus Close To Job & Lock It |  | Maintain Distance Between Job & Stylus
|  |  |  | 7 | Double Click On 'Cam Tester' Icon, Click & Select Work No. & Name as the case may be i.e. Pin Origin-KSP, No.-1 Cam OriginKSP,Pin Origin KRB |  | Work No. & Name Selection To Be Done Carefully
Cam tester software icons |  | program details feed in software |  | 8 | Click 'MAKE NOMINAL' , Check 'Cam Position' & Then Press'End' |  | Check Carriage\Stylus Position by moving follwer carraige manually to the cam piece and check the feeded value of carraige position for both cam piece In and Ex are same.
do not put your hand into the moving parts during operation
|  |  |  | 9 | Now Push Carriage Position Button Left/Right To Verify Cam Position
|  |  |  | 10 | In Case Of Pin Origin Place Slip(Of Thickness Equal To Radius Of LockPin) On Height Master Gauge & With Help Of Electronic Dial Unit Make High Point Of Pin & Height Master Gauge At Same Level
Measuring report screen |  | Adjustment screen
|  |  |  | 11 | In Case Of No-1 Cam Origin, Make Lift & Angle Zero At Highest Point Of EX Cam Piece
|  |  |  | 12 | Now Go To 'Measurment Mode', Click 'Measure And Report' Then 'END' & Finally Click 'START' Serially, to Initiate Measurment
|  |  |  | 13 | Push 'Auto' Button, Then 'Measure' & Carriage Will Move To Cam After Reaching Home Position
Knock Pin & Tightning Screw Opposite To Each Other |  | Electronic Display Unit |  | 14 | Put Paper On Plotter & Press Chart Hold Button
|  |  |  | 15 | After Finishing Measurment Carriage Will Come To Home Position . Printer & Plotter Will Start Printing |  | During Home Position Setup Probe Should not Hit Job
|  |  |  | 16 | Push Manual Button, Using Sliding Screw Move Probe Away From Job, Loose Locking Srews & Unclamp The Job
17
SAFETY DEVICES TO BE USED |  |  |  | ATTENTION |  |  |  |  | WORKING HEADS
Name of safety device |  | Qty | Remarks | Route to be followed in case of any trouble |  |  |  |  | 1. MACHINE STARTUP |  | 5. MODEL CHANGE
[9 empty] | 2. MACHINE OPERATION |  | 6. MACHINE STOP
|  |  |  | A Shift, General Shift |  | Line Asso. Exec. Sr. Exe. A.M D.M.
[9 empty] | 3. QUALITY MEASUREMENT |  | 7. CLEANING
|  |  |  | B & C Shift |  | Line Asso. Exec. Sr. Exe. | [6 empty] | VITAL CHECK POINTS ARE SHOWN IN RED COLOUR
4. TOOL CHANGE

## CamProfileTester 8.5.2017
[7 empty] | OPERATION STANDARD | [5 empty] | CONTROL NUMBER SOP-QA-40
Equipment Name |  | Cam Tester |  |  | Process | CAM PROFILE & ANGLE MEASUREMENT |  | Section | Date | 8.5.2017 |  | PREPARED |  | CHECKED | APPROVED
Components |  | ALL Cam shaft |  |  | Oprn. No. |  |  | Quality Assurance | Issue No | 02 |  | Shivananda |  | Prashanth | Sathoshi san
[9 empty] | Rev.NO | 00
|  |  |  | SI.No. | Operation Procedure |  | Vital Point Of Operation |  | Things which You should Not Do |  |  | Result in case you do it
MAIN MACHINE
|  |  |  | 1 | Switch On UPS
|  |  |  | 2 | Switch On Main Power Supply , CPU & MASTER ON Cam Tester M/c Sequentially |  | Check For Power Supply |  | Don't Run Computer On UPS |  |  | UPS May Get Trip
|  |  |  | 3 | Clean the Head stock center & Tail stock center by using fine cloth |  | check condition of the centers
|  |  |  | 4 | install the component to the work driving dog rotator |  | Maintain Distance Between Stylus & Driving Jig |  | Stylus Should Not Hit By Job In Any Case |  |  | Stylus May Get Damage
Slightly Tight Job In Driven Jig
|  |  |  | 5 | Move the tailstock to a point where it can clamp the workpiece, lock the tailstock. |  | In Case Of Shaft Comp Cam Lock Pin Must Be Opposite To Locking Sceews Of Driven Jig
Maintain optimum distance between driven jig and Stylus to avoid any damage to follower/stylus.
|  |  |  | 6 | Unlock Sliding Screw, Move Stylus Close To Job & Lock It |  | Maintain Distance Between Job & Stylus
|  |  |  | 7 | Double Click On 'Cam Tester' Icon, Click & Select Work No. & Name as the case may be i.e. Pin Origin-KSP, No.-1 Cam OriginKSP,Pin Origin KRB |  | Work No. & Name Selection To Be Done Carefully
Cam tester software icons |  | program details feed in software |  | 8 | Click 'MAKE NOMINAL' , Check 'Cam Position' & Then Press'End' |  | Check Carriage\Stylus Position by moving follwer carraige manually to the cam piece and check the feeded value of carraige position for both cam piece In and Ex are same.
do not put your hand into the moving parts during operation
|  |  |  | 9 | Now Push Carriage Position Button Left/Right To Verify Cam Position
|  |  |  | 10 | In Case Of Pin Origin Place Slip(Of Thickness Equal To Radius Of LockPin) On Height Master Gauge & With Help Of Electronic Dial Unit Make High Point Of Pin & Height Master Gauge At Same Level
Measuring report screen |  | Adjustment screen
|  |  |  | 11 | In Case Of No-1 Cam Origin, Make Lift & Angle Zero At Highest Point Of EX Cam Piece
|  |  |  | 12 | Now Go To 'Measurment Mode', Click 'Measure And Report' Then 'END' & Finally Click 'START' Serially, to Initiate Measurment
|  |  |  | 13 | Push 'Auto' Button, Then 'Measure' & Carriage Will Move To Cam After Reaching Home Position
Knock Pin & Tightning Screw Opposite To Each Other |  | Electronic Display Unit |  | 14 | Put Paper On Plotter & Press Chart Hold Button
|  |  |  | 15 | After Finishing Measurment Carriage Will Come To Home Position . Printer & Plotter Will Start Printing |  | During Home Position Setup Probe Should not Hit Job
|  |  |  | 16 | Push Manual Button, Using Sliding Screw Move Probe Away From Job, Loose Locking Srews & Unclamp The Job
17
SAFETY DEVICES TO BE USED |  |  |  | ATTENTION |  |  |  |  | WORKING HEADS
Name of safety device |  | Qty | Remarks | Route to be followed in case of any trouble |  |  |  |  | 1. MACHINE STARTUP |  | 5. MODEL CHANGE
[9 empty] | 2. MACHINE OPERATION |  | 6. MACHINE STOP
|  |  |  | A Shift, General Shift |  | Line Asso. Exec. Sr. Exe. A.M D.M.
[9 empty] | 3. QUALITY MEASUREMENT |  | 7. CLEANING
|  |  |  | B & C Shift |  | Line Asso. Exec. Sr. Exe. | [6 empty] | VITAL CHECK POINTS ARE SHOWN IN RED COLOUR
4. TOOL CHANGE

## CamProfileTester 5.1.2019
[7 empty] | OPERATION STANDARD | [5 empty] | CONTROL NUMBER SOP-QA-40
Equipment Name |  | Cam Tester |  |  | Process | CAM PROFILE & ANGLE MEASUREMENT |  | Section | Date | 5.1.2019 |  | PREPARED |  | CHECKED | APPROVED
Components |  | ALL Cam shaft |  |  | Oprn. No. |  |  | Quality Assurance | Issue No | 02 |  | Shivananda |  | Prashanth | Umasankar
[9 empty] | Rev.NO | 01
|  |  |  | SI.No. | Operation Procedure |  | Vital Point Of Operation |  | Things which You should Not Do |  |  | Result in case you do it
MAIN MACHINE
|  |  |  | 1 | Switch On UPS
|  |  |  | 2 | Switch On Main Power Supply , CPU & MASTER ON Cam Tester M/c Sequentially |  | Check For Power Supply |  | Don't Run Computer On UPS |  |  | UPS May Get Trip
|  |  |  | 3 | Clean the Head stock center & Tail stock center by using fine cloth |  | check condition of the centers
|  |  |  | 4 | install the component to the work driving dog rotator |  | Maintain Distance Between Stylus & Driving Jig |  | Stylus Should Not Hit By Job In Any Case |  |  | Stylus May Get Damage
Slightly Tight Job In Driven Jig
|  |  |  | 5 | Move the tailstock to a point where it can clamp the workpiece, lock the tailstock. |  | In Case Of Shaft Comp Cam Lock Pin Must Be Opposite To Locking Sceews Of Driven Jig
Maintain optimum distance between driven jig and Stylus to avoid any damage to follower/stylus.
|  |  |  | 6 | Unlock Sliding Screw, Move Stylus Close To Job & Lock It |  | Maintain Distance Between Job & Stylus
|  |  |  | 7 | Double Click On 'Cam Tester' Icon, Click & Select Work No. & Name as the case may be i.e. Pin Origin-KSP, No.-1 Cam OriginKSP,Pin Origin KRB |  | Work No. & Name Selection To Be Done Carefully
Cam tester software icons |  | program details feed in software
|  |  |  | 8 | Click 'MAKE NOMINAL' , Check 'Cam Position' & Then Press'End' |  | Check Carriage\Stylus Position by moving follwer carraige manually to the cam piece and check the feeded value of carraige position for both cam piece In and Ex are same.
do not put your hand into the moving parts during operation
|  |  |  | 9 | Now Push Carriage Position Button Left/Right To Verify Cam Position
|  |  |  | 10 | In Case Of Pin Origin Place Slip(Of Thickness Equal To Radius Of LockPin) On Height Master Gauge & With Help Of Electronic Dial Unit Make High Point Of Pin & Height Master Gauge At Same Level
Measuring report screen |  | Adjustment screen
|  |  |  | 11 | In Case Of No-1 Cam Origin, Make Lift & Angle Zero At Highest Point Of EX Cam Piece
|  |  |  | 12 | Now Go To 'Measurment Mode', Click 'Measure And Report' Then 'END' & Finally Click 'START' Serially, to Initiate Measurment
Knock Pin & Tightning Screw Opposite To Each Other |  | Electronic Display Unit
|  |  |  | 13 | Push 'Auto' Button, Then 'Measure' & Carriage Will Move To Cam After Reaching Home Position
|  |  |  | 14 | Put Paper On Plotter & Press Chart Hold Button
|  |  |  | 15 | After Finishing Measurment Carriage Will Come To Home Position . Printer & Plotter Will Start Printing |  | During Home Position Setup Probe Should not Hit Job
|  |  |  | 16 | Push Manual Button, Using Sliding Screw Move Probe Away From Job, Loose Locking Srews & Unclamp The Job
|  |  |  | 17 | for checking the symmetrycity put the slip guage into keyway groove and make alignment zero by using height guage |  | Check the symmetricity
|  |  |  | 18 | after finishing alignment take the high point of the bearing OD , note the reading on height guage display and repeat the same procedure in the opposite side and note down the reading
symmetricity inspection setting
|  |  |  | 19 | taking the diffrence of the both readings and dividing the diffrence reading by 2 calculate the symmetricity.
SAFETY DEVICES TO BE USED |  |  |  | ATTENTION |  |  |  |  | WORKING HEADS
Name of safety device |  | Qty | Remarks | Route to be followed in case of any trouble |  |  |  |  | 1. MACHINE STARTUP |  | 5. MODEL CHANGE
[9 empty] | 2. MACHINE OPERATION |  | 6. MACHINE STOP
|  |  |  | A Shift, General Shift |  | Line Asso. Exec. Sr. Exe. A.M D.M.
[9 empty] | 3. QUALITY MEASUREMENT |  | 7. CLEANING
|  |  |  | B & C Shift |  | Line Asso. Exec. Sr. Exe. | [6 empty] | VITAL CHECK POINTS ARE SHOWN IN RED COLOUR
4. TOOL CHANGE