## Cover Page
| Doc No : | ISD-QA-008
Musashi Auto Parts India Pvt Ltd | Issue No : | 6
| Rev No : | 0
| Date : | 01.04.2015
LAB MANUAL
(Standard Room)
Checked By
Prepared By
Released By
Approved By

## Ownership
| Doc No : | ISD-QA-008A
Musashi Auto Parts India Pvt Ltd | Issue No : | 6
| Rev No : | 0
| Date : | 01.04.2015
Lab Manual (Standard Room)
This Lab Manual (Standard Room) is the property of
Musashi Auto Parts India Pvt Limited and shall not be
copied or reproduced without the written consent
of the Works Director or Works Head.
Address:
M/s Musashi Auto Parts India Pvt Ltd
Plot No 33-35 & 46-60
Sector 7,Bawal Industrial Estate,
Bawal District,Rewari,Haryana
India
Checked By
Prepared By
Released By
Approved By

## Contents Table
|  | Doc No : | ISD-QA-008B
Musashi Auto Parts India Pvt Ltd |  | Issue No : | 6
|  | Rev No : | 0
|  | Date : | 01.04.2015
Table of Contents
1 | Scope.
2 | Associates/Line Associates & their Educational Qualification.
3 | Skill Matrix.
4 | List of Measuring Instruments.
5 | Enviromental Conditions.
6 | Measuring Instruments & operation Standards
Checked By
Prepared By
Released By
Approved By

## Scope
| Doc No : | ISD-QA-008C
Musashi Auto Parts India Pvt Ltd | Issue No : | 6
| Rev No : | 0
| Date : | 01.04.2015
Scope
This manual aims to describe the documented description/procedures of
entire range of activities carried out,facility available for inspection ,test
and calibration of various measuring instruments available in Standard
Room.
Checked By
Prepared By
Released By
Approved By

## Qualification-p
|  |  | Doc No : |  | ISD-QA-008D
Musashi Auto Parts India Pvt Ltd |  |  | Issue No : |  | 6
|  |  | Rev No : |  | 0
|  |  | Date : |  | 01.04.2015
Manpower Avaialble/Educational Qualification & Designation
Sr No | Name | Qualification |  | Experience (in years) | Designation
1
2
3
4
5
6
7
8
9
10
11
12
13
Checked By
Prepared By
Released By
Approved By

## Competancy Reqd
Annexure to Laboratory Scope
Competancy Reqd
Purpose : The intent of this chart is to provide the guidelines for who all line associates are authorised to operate on individual
machines stationed in Standard Room.
Equpment Name
CMM | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of MCOSMOS software & various functions.
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Contour Tester | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of Formtracepak software & various functions.
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Profile Projector | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of mannual inspection of dia,angle & distance.
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Cam Tester | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of Cam tester software & various functions.
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Cam Tester (Adcole) | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of Adcole911 software & various functions.
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Barkhosen Testing Machine | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of viewscan software and various functions etc..
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Roughness Tester (Mitutoyo) | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of roughness software and various functions etc..
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Kosaka Roughness Tetsr | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of SGA31 software, various function & knowledge of roughness perameters.
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
XRF Machine | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of XRF software & various function.
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Roundness Tetser | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of formpack software & form errors viz Roundness,Taper,Ovality,Cylindricity error etc..
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
CNC Gear Checker | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of Gear inspection parameters viz Lead Error,Profile Error,Crowning error etc..
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Optical Shaft Measuring Machine | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of jenoptic software and various functions etc..
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Linear Height Gauge | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Bore Scope | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings
Calibration Instruments (Electronic Dial Calibration Tester / Caliper Checker ) | Has worked under the supervision of immediate incharge/skilled person for atleast one week time
Basic understanding of calibrations methods etc..
Able to evaluate the results so obtained
Basic understanding of process sheet/engineering drawings

## Competancy Reqd (2)
Authorised to work on the machine: Yes |  |  |  |  | Not authorised to work on the machine : No
Line Associates Name |  |  | CMM | Contour Tester | Profile Projector | Cam Tester | Cam Tester (Adcole) | Stresstech | Roughness Tester (Mitutoyo) | Kosaka Roughness Tetsr | XRF Machine | Roundness Tester | CNC Gear Tester | Optical Shaft Measurment | Linear Height Gauge | Bore Scope | Calibration Instruments (Electronic Dial Calibration Tester / Caliper Checker )
KULDEEP SINGH |  |  | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO
KANHIYA | KANHIYA | KANHIYA | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO
SANTOSH KUMAR | SANTOSH KUMAR | SANTOSH KUMAR | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO
MAMTA SINGH | MAMTA SINGH | MAMTA SINGH | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | NO | NO | YES | NO | NO
RAVINDER GHANGAS | RAVINDER GHANGAS | RAVINDER GHANGAS | NO | NO | NO | YES | NO | NO | YES | NO | NO | YES | YES | NO | YES | NO | NO
UMESH | UMESH | UMESH | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | YES | YES | YES | NO | NO
SACHIN KUMAR | SACHIN KUMAR | SACHIN KUMAR | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO
RAJAT | RAJAT | RAJAT | NO | YES | YES | NO | YES | YES | YES | NO | YES | YES | YES | YES | YES | YES | NO
ABHAY | ABHAY | ABHAY | NO | YES | YES | YES | YES | YES | YES | NO | YES | YES | YES | NO | YES | YES | NO
HEMANT | HEMANT | HEMANT | NO | YES | NO | YES | YES | YES | YES | NO | YES | YES | YES | NO | YES | YES | NO
VIKAS | VIKAS | VIKAS | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | YES | NO
AMIT KUMAR | AMIT KUMAR | AMIT KUMAR | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO
BASANT | BASANT | BASANT | NO | YES | NO | NO | YES | YES | YES | NO | NO | YES | YES | NO | YES | YES | NO
SUSHANT | SUSHANT | SUSHANT | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | YES | YES | YES | YES | NO
BILAL AHAMD | BILAL AHAMD | BILAL AHAMD | NO | YES | NO | NO | YES | YES | YES | NO | NO | NO | YES | NO | YES | YES | NO
ROBIN | ROBIN | ROBIN | NO | YES | YES | NO | YES | YES | YES | YES | NO | YES | YES | NO | YES | YES | NO
ALWAN HASSAN | ALWAN HASSAN | ALWAN HASSAN | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | NO
VINAY SINGH | VINAY SINGH | VINAY SINGH | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | NO
SANDEEP | SANDEEP | SANDEEP | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | NO
AJEET KUMAR | AJEET KUMAR | AJEET KUMAR | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | NO
KARAN | KARAN | KARAN | NO | NO | NO | NO | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | NO
RANI | RANI | RANI | NO | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO | NO | YES | YES | NO
JYOTI YADAV | JYOTI YADAV | JYOTI YADAV | NO | NO | NO | NO | YES | YES | YES | YES | YES | YES | NO | NO | YES | YES | NO
ARTI SHARMA | ARTI SHARMA | ARTI SHARMA | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | NO
ANUPAM YADAV | ANUPAM YADAV | ANUPAM YADAV | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | NO
POONAM DEVI | POONAM DEVI | POONAM DEVI | NO | NO | NO | NO | YES | YES | YES | NO | NO | NO | YES | NO | YES | YES | NO
AKSAH MITTAL | AKSAH MITTAL | AKSAH MITTAL | NO | YES | YES | NO | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO
ANUJ PATEL | ANUJ PATEL | ANUJ PATEL | NO | YES | YES | YES | YES | YES | YES | NO | YES | YES | YES | NO | YES | YES | NO
ARUN MOHANTA | ARUN MOHANTA | ARUN MOHANTA | NO | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO
SACHIN CHHILLAR | SACHIN CHHILLAR | SACHIN CHHILLAR | NO | NO | NO | NO | YES | YES | YES | NO | NO | YES | NO | NO | YES | NO | NO
ASHUTOSH | ASHUTOSH | ASHUTOSH | NO | NO | NO | NO | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | NO
RAJAN | RAJAN | RAJAN | NO | YES | YES | NO | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | NO
MUKESH KUMAR | MUKESH KUMAR | MUKESH KUMAR | NO | YES | YES | NO | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | NO
ARVIND | ARVIND | ARVIND | NO | YES | NO | NO | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | NO
HITESH KUMAR | HITESH KUMAR | HITESH KUMAR | NO | NO | NO | NO | YES | YES | YES | NO | NO | NO | YES | NO | YES | NO | NO
PREETI | PREETI | PREETI | YES | YES | NO | YES | NO | NO | YES | NO | NO | YES | YES | NO | YES | YES | YES
SATANDAR SHARMA | SATANDAR SHARMA | SATANDAR SHARMA | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | YES
SHUBHAM | SHUBHAM | SHUBHAM | YES | YES | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | YES
VINITA PATHERA | VINITA PATHERA | VINITA PATHERA | NO | YES | NO | NO | NO | NO | YES | NO | NO | NO | NO | NO | YES | NO | YES
NEELAMBAR MISHRA | NEELAMBAR MISHRA | NEELAMBAR MISHRA | YES | YES | YES | YES | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | YES
JANGBHADUR | JANGBHADUR | JANGBHADUR | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | YES
ATUL KUMAR TIWARI | ATUL KUMAR TIWARI | ATUL KUMAR TIWARI | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | YES
AWNEESH | AWNEESH | AWNEESH | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | YES | NO | YES | NO | YES
TAAPAS | TAPAS | TAPAS | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | YES
ARVIND | ARVIND | ARVIND | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | YES
JITENDRA | JITENDRA | JITENDRA | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | YES
RINKU | RINKU | RINKU | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | YES
PRAVIND | PRAVIND | PRAVIND | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | NO | YES | NO | YES
PRAVEEN | PRAVEEN | PRAVEEN | NO | NO | NO | NO | NO | NO | YES | NO | NO | NO | YES | NO | YES | NO | YES
RAMSUNDER-2W |  |  | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | NO
SANJAY-4W |  |  | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | YES | NO | YES | YES | NO

## Skill Matrix
| Doc No : | ISD-QA-008E
| Issue No : | 6
Musashi Auto Parts India Pvt Ltd | Rev No : | 0
| Date : | 01.04.2015
Skill Matrix
The Skill Matrix of the all the line associates are described in the list
attached.The whole skill matrix level of various line associates are
constantly upgraded on quartely basis to ensure them the continous
learning.
Enclosed: Skill Matrix Sheet
Checked By
Prepared By
Released By
Approved By

## Annexure to laboratory scope
Annexure to Laboratory Scope
Sr No | Equipments | Equipment Model | Standards | Methods/Work Instruction
1 | Gear Profile Tester | Osaka GC-3H | As per the guidelines given in Process Sheet/Final Drawing | Annexure to ISW-QA-001
2 | Roughness Tester | Mitutoyo SJ 400 | As per the guidelines given in Process Sheet/Final Drawing | Annexure to ISW-QA-004
3 | Linear Height Gauge | Mitutoyo LH 600 | As per the guidelines given in Process Sheet/Final Drawing | Annexure to ISW-QA-005
4 | Electronic Dial Calibration Tester | Kudale | JIS B 7503 | Annexure to ISW-QA-006
5 | Outside Micrometer | -- | JIS B 7502 | Annexure to ISW-QA-007
6 | Height Gauge | -- | JIS B 7517 | Annexure to ISW-QA-008
7 | Plunger Dial Gauge | -- | JIS B 7503 | ISW-QA-009
8 | Vernier Caliper | -- | JIS B 7507 | ISW-QA-010
9 | Bench Center | -- | IS 5980-1978 | ISW-QA-011
10 | Bevel Protractor | -- | IS 4239-1970 | ISW-QA-012
11 | Cylendrical Pins & Ball Anvils | -- | IS 11103-1984 | ISW-QA-013
12 | Taper Mandrel | -- | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-014
13 | Dial Test Indicator | -- | JIS B 7533 | ISW-QA-015
14 | Bore Gauge | -- | JIS B 7515 | ISW-QA-016
15 | Spline Plug Gauge | -- | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-017
16 | Spline Ring Gauge | -- | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-018
17 | Disc Micrometer | -- | JIS B 7502 | ISW-QA-019
18 | Height Master Gauge | -- | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-020
19 | Plain Plug Gauge | -- | JIS B 7420:1997 | ISW-QA-021
20 | Snap Gauge | -- | JIS B 7420:1997 | ISW-QA-022
21 | Plain Ring Gauge | -- | JIS B 7420:1997 | ISW-QA-023
22 | Width Gauge | -- | JIS B 7420:1997 | ISW-QA-024
23 | Feeler Gauge | -- | JIS B 7524:1992 | ISW-QA-025
24 | Try Square | -- | JIS B 7526:1995 | ISW-QA-026
25 | Steel Parallel | -- | JIS B 7514:1977 | ISW-QA-027
26 | Magnetic V Block | -- | JIS B 7540:1972 | ISW-QA-028
27 | Air Plug Gauge | -- | As per maker catalogue | ISW-QA-029
28 | Cam Tester | Accretech | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-030
29 | CNC Gear Checker | CLP-35 | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-031
Sr No | Equipments | Equipment Model | Standards | Methods/Work Instruction
30 | Coordinate Meas Machine | Mitutoyo-574 | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-032
31 | Roundness Tester (RA-2000) | RA-2200 | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-033
32 | Contour Tester (CV-3000) | CV-3000 | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-034
33 | Kosaka Roughness Tester | -- | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-035
34 | Profile Projector | Nikkon V12B | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-036
35 | Cam Tester | ADCOLE-911 | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-037
36 | Contour Tester (CV-3100) |  | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-038
37 | Contour Tester (CV-3200) |  | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-039
38 | Roundness Tester(RA-2200) |  | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-040
39 | Roundness Tester(Kosaka) |  | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-041
40 | Bore Scope |  | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-042
41 | Barkhousan Testing Machine | -- | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-043
42 | Optical Shaft Measurment. | -- | As per the guidelines given in Process Sheet/Final Drawing | ISW-QA-044
43 | Master Assembly | -- | Gauge Drawings | ISW-QA-045
44 | Surface Plate | -- | JIS B 7513 | ISW-QA-046
45 | Master Gear | -- | Gauge Drawings | ISW-QA-047
46 | Blade Micrometer | -- | JIS B 7502 | ISW-QA-048
47 | Gear Tooth Micrometer | -- | JIS B 7502 | ISW-QA-049
48 | Point Micrometer | -- | JIS B 7502 | ISW-QA-050
49 | Depth Micrometer | -- | JIS B 7544 | ISW-QA-051
50 | Inside Groove Vernier | -- | JIS B 7507 | ISW-QA-052
51 | Depth Vernier | -- | JIS B 7518 | ISW-QA-053
52 | Pasa Meter | -- | JIS B 7520 | ISW-QA-054
53 | PCD Tester | -- | IS 5980-1978 | ISW-QA-055
54 | Slip Gauge Block | -- | JIS B 7506 : 1997 | ISW-QA-056
55 | Calliper Checker | -- | As per maker catalogue | ISW-QA-057
56 | Micrometer Checker | -- | As per maker catalogue | ISW-QA-058
57 | Straight Mandrel | -- | JIS B 7545 : 1982 | ISW-QA-059
58 | Air Converter | -- | As per maker catalogue | ISW-QA-060
59 | Air Gauge Unit | -- | As per maker catalogue | ISW-QA-061
Sr No | Equipments | Equipment Model | Standards | Methods/Work Instruction
60 | Column Bar | -- | As per maker catalogue | ISW-QA-062
61 | Comparator Stand | -- | As per maker catalogue | ISW-QA-063
62 | Load Cell | -- | As per maker catalogue | ISW-QA-064
63 | Vertical Gear Roll Tester | -- | As per maker catalogue | ISW-QA-065
64 | Horizontal Gear Roll Tester | -- | As per maker catalogue | ISW-QA-066
65 | Bend & Twist FIxture | -- | As per maker catalogue | ISW-QA-067
66 | Functional Gauges | -- | Gauge Drawings | ISW-QA-068
67 | Mic Setting Master | -- | As per maker catalogue | ISW-QA-069
68 | Span Master | -- | Gauge Drawings | ISW-QA-070
69 | Pasameter Master | -- | Gauge Drawings | ISW-QA-071
70 | Spline Arbor | -- | Gauge Drawings | ISW-QA-072
71 | Master Disc | -- | Gauge Drawings | ISW-QA-073
72 | Setting Ring | -- | As per maker catalogue | ISW-QA-074
73 | Thread Ring Gauge | -- | JIS B 0251:1998 | ISW-QA-075
74 | Thread Plug Gauge | -- | JIS B 0251:1998 | ISW-QA-076
75 | Air Snap Gauge | -- | As per maker catalogue | ISW-QA-077
76 | Air Ring Gauge | -- | As per maker catalogue | ISW-QA-078
77 | Dial Snap Gauge | -- | JIS B 7520 | ISW-QA-079
78 | Magnetic Dial Stand | -- | As per maker catalogue | ISW-QA-080
79 | Taper Cone Gauges | -- | Gauge Drawings | ISW-QA-081
80 | GRT Nick NG Masters | -- | Process Sheet | ISW-QA-082

## Ext Agency-p
Annexure to Laboratory Scope
LIST OF EXTERNAL LABS/AGENCIES APPROVED FOR JOB WORKS
SRL | AGENCY/LAB | CRITERIA FOR SELECTION | JOBS TO BE CARRIED OUT
1 | A.V Industries | NABL Approved | Calibration of Bench center, Surface Plate,Pressure Gauge Blocks Temp Controller etc.
2 | Osaka Simetsu | Osaka Machine Manufacturer | Calibration of CNC Gear Tester & Manual Gear Tester Osaka Make.
3 | Creative Calibration Center | NABL Approved | Calibration of WCP & WCR etc.
4 | Baker Gauges | Traceability of Master to NPL | Calibration of Air units
5 | Mitutoyo South Asia | Traceability of Master to NPL | Calibration of Contour tester, Roundness tester, Cmm & Liner height gauge.
6 | Mikronix Gauges | NABL Approved | Calibration of Gauge Block.
7 | Precision Product | Kosaka Approved | Calibration of Roughness tester & Roundness tester.
8 | Adcole | Adcole Machine Manufacturer | Calibration of Cam Tester Adcole 911
9 | Roll Scan 250 | Stresstech | Calibration of Stresstech (Roll Scan 250)
10 | I R Technology services pvt.ltd. | OLYMPUS Approved | Calibration of XRF Material Analyzer
11 | Optical Shaft Measurment | Jenoptic | Calibration of Optical Shaft Measurment.
12 | Tosai engineering pvt. ltd. | Accretech | Calibration of Cam Tester Accretech.
| Prepared On | 06.10.2010
| Approved On | 07.10.2010
| Reviewed On | 15.12.2011
| Next Review On | 16.12.2012
| Prepared By | Checked By |  | Approved By
REV 01

## Conditions
|  |  | Doc No : | ISD-QA-008F
|  |  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : | 0
|  |  | Date : | 01.04.2015
Conditions to be considered for Standard Room
1 | Temperature | : | 20 deg ± 1 deg C
2 | Humidity | : | No Excessive humidity
3 | Insulation | : | --
4 | Vibration | : | +ve/-ve vibrations
5 | Furniture | : | Steel or Wood (White or Light)
6 | Electrical Voltage Regulation | : | Stablised & Controlled Voltage
7 | Demagnetisation | : | First Demagnetise
8 | Reference Standard | : | Proper Traceability
9 | Calibration Equipment | : | Proper Traceability
10 | Temperature Stablisation | : | 1 hrs for normal
|  | : | 2 hrs for Master
11 | Entry Restricted
Checked By
Prepared By
Released By
Approved By

## Facilities
|  |  | Doc No : |  | ISD-QA-008G
|  |  | Issue No : |  | 4
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : |  | 2
|  |  | Date : |  | 12.08.2014
Inspection/Calibration Facility Avaialble
The various Inspection/Calibration Facilities available in the Lab are :
Sr No | Instrument Name | Make |  | Model | Installation (Month/Year)
1 | Gear Profile Tester | Osaka |  | GC - 3H | Oct '03
2 | CNC Gear Checker-I | Osaka |  | CLP 35 | April '05
3 | CNC Gear Checker-II | Osaka |  | CLP 35 | Dec '08
4 | CNC Gear Checker-III | Osaka |  | CLP 35 | Nov'11
5 | CNC Gear Checker-IV | Osaka |  | CLP 35 | Sep'12
6 | CNC Gear Checker-V | Osaka |  | CLP 35 | June'13
7 | CNC Gear Checker-VI | Osaka |  | CLP 35 | July '14
8 | CNC Gear Checker-VII | Osaka |  | CLP 35S | July '16
9 | Roundness Tester | Mitutoyo |  | RA 2000 DH | Oct '03
10 | Roundness Tester | Mitutoyo |  | RA 2200 DH | Nov'12
11 | Roundness Tester | Kosaka |  | EC 2500 | Nov'11
12 | Roughness Tester | Mitutoyo |  | SJ 402 | Oct'03
13 | Contour Tester | Mitutoyo |  | CV 3000 | Feb'12
14 | Contour Tester | Mitutoyo |  | CV 3100 | Feb'12
15 | Contour Tester | Mitutoyo |  | CV 3200 | Nov'12
Sr No | Instrument Name | Make |  | Model | Installation (Month/Year)
16 | Dial Calibration Tester | Kudale Instruments |  | ---- | Sep'03
17 | Caliper Checker | Mikronix |  | ---- | Sep'03
18 | Linear Height Gauge-1 | Mitutoyo |  | LH600 | Oct'03
19 | Linear Height Gauge-2 | Mitutoyo |  | LH600 | Feb'12
20 | Linear Height Gauge-3 | Mitutoyo |  | LH600 | Nov'12
21 | Cam Tester | Tosei Engineering Corp. |  | ---- | Oct'03
22 | Cam Tester | Adcole |  | 911 | Feb'12
23 | Coordinate Measuring Machine | Mitutoyo |  | Crysta 574 | June'05
24 | Roughness Tester (Bibri Checking) | Kosaka |  | SE-3500 | July '07
25 | Roughness Tester (Bibri Checking) | Kosaka |  | SE-3500 | Feb'12
26 | Profile Projector | Nikkon |  | V12B | Aug'11
27 | Barkhousen Testing M/c | Stresstech |  | Cam Scan 300 | Jan'12
28 | Optical Shaft Masuring M/c | Janoptic |  | ----- | Jan'13
29 | Bore Scope | Olympus |  | ----- | March'12
Checked By
Prepared By
Released By
Approved By

## Gear Profile Tester
|  |  |  | Doc No : |  | ISW-QA-001
|  |  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : |  | 0
|  |  |  | Date : |  | 01.04.2015
Gear Profile Tester (GC-3H)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of Gear Profile/Lead for all types of spur/helical Gears and Shafts in the capacity of
max.outside Dia 350 mm with centers and max.shaft length 30~ 400 mm.
Reference Standards/Materials :
Master Involute & Lead Test Mandrel (Straight)
Description: | This machine makes use of a base circle disc attached with the lower center.The gear is set in
between the upper and lower center .The lower center is fixed using the dog and the carrier.
The base circle disc is in brought in contact with the straight edge by manually feeding the gear
support towards left direction .By rotary motion of the base circle disc and the longitudinal
movement of the straight edge ,the involute curve is generated between stylus and the tooth flank
lead is accomplished by moving the measuring stylus/post in veritcal direction .The straight edge
longitudinally gets shifted in accordance with the vertical motion of the measuring post and the
lead is generated between the stylus and tooth flank of the gear to be measured.The lead errors
are recorded on recording papers.
Standard Accessories :
The standard accessries required for the operation of the machine are Stylus,Lower Center,Dog and
carrier,recording ink,recording pen and the chart paper.
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.In case
of any doubt in the overall performance of the machine,the periodic calibration is done using a Test
Mandrel and Master Involute & Lead supplied along with the machine by the maker .
The master Involute and Lead is loaded between centers and lead and profile for different helix
angles installation.
Calibration Frequecy:
Once in a Year & In case of suspected performance, the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-001"
Periodic Calibration of Gear Profile Tester (GC-3H) - "ISW-QA-001A"
page 1 of 2
Permissible Deviations for the Gear Profile Tester (GC-3H) :
Sr No | Test Characteristics | Test Method |  |  | Acceptable Limits of Error (mm)
1 | Runout on Test Bar | Measuring by the putting the Test Bar between upper & lower center and clamp properly. Rotate the Test Bar & take the observations on different places on Test Bar. |  |  | Runout Should be with in 0.002
2 | Parallelism on Test Bar | Measuring by the putting the Test Bar between upper & lower center and clamp properly. Move the measuring slide & take the observations on different places on Test Bar. |  | Runout Should be with in 0.002 | Runout Should be with in 0.002
3 | Difference of outernal Dimeter on test bar | Measuring by the putting the test Bar between upper & lower center and clamp properly.Rotate the Test Bar & take the observations on difference of outernal Diameter Between A & B At the Point of (2) |  |  | Difference between A & B Shoud be With in 0.002
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | OSAKA SEIMITSU | CNC Gear Tester
Checked By
Prepared By
Released By
Approved By
Checked By
Prepared By
Released By
Approved By

## Roughness Tester
|  |  |  | Doc No : | ISW-QA-004
|  |  |  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.04.2015
Roughness Tester (SJ-402)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of various roughness parameters of various machines components
Reference Standards/Materials :
1) Levelling Table
2) Judgment doc. No=178-602
3) Precision Reference Specimen
Code No:178-602 A & 300210
Description: | This machine makes use of detector with exchangeable nosepiece that is interchangeable between
skidless and skid measurement.In skidless measurement,surface irregularities are detected
with reference to guide on the drive unit.In skid measurement surface irregularities are detected with
reference to the skid that traces waviness on a measuring surface.The stylus is positioned on the
component perpendicular to the machining direction parallel to the detector and respective roughness
parameters are evaluated.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus,
(Part No 12AAB403 5 microns tip radius),Sine Vise,XY Levelling Table with swivel etc….
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done.
The periodic calibration is done using Precision Reference Specimen described above.The reference
specimen is placed on the levelling table and the table is levelled using the options available.The
stylus is made to traverse a decided cut off length and roughness value obserbed is compared with
the master value specified on the precision specimen.In case of any mismatch the conversion factor is
applied and the machine is recalibrated using the master specimen.
Calibration Frequecy:
Once in a Year & In case of suspected performance, the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-004"
Periodic Calibration of Roughness Tester (SJ-402) - "ISW-QA-004A"
page 1 of 2
Permissible Deviations for the Roughness Specimen :
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (µm)
1 | Check Surface Finish of Specimen in RA | Measuring the Specimen first adjust Parallelism between the drive unit and the specimen and then measure the surface finish. Measuring Condition :- Cut-off lenght - 2.5mm Evaluation lenght -12.5mm Measuring point - 5 Points Filter - PC 50% |  | Surface Finish Should be with in ± 3% of nominal value.
2 | Check Surface Finish of Specimen in RY |  |  | Surface Finish Should be with in ± 4% of nominal value.
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | Mitutoyo | Roughness Specimen Master
Checked By
Prepared By
Released By
Approved By
Roughness Tester (SJ-402)
Checked By
Prepared By
Released By
Approved By

## Linear Height Gauge
|  |  | Doc No : |  | ISW-QA-005
|  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : |  | 0
|  |  | Date : |  | 01.04.2015
Linear Height Gauge (LH-600B)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of various linear 1D & 2 D dimensional parameters viz circle diameter,distance etc.
of machined components
Reference Standards/Materials :
1) Setting Block
Description: | This machine could be used effectively for measurement of one dimensional and two dimensional
linear dimensions with great accuracy and repeatability.This measuring instrument could be used
to measure linear dimensions in the capacity of 600 mm Max length.Before measurement ,the
instrument needs to understand its absolute origin and should sense the compensation reference
as well.The sensing of compensation reference and absolute origin is done with the help of Setting
block provided with the machine.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus, Probe Diameter
Calibration Block (12AAA787) etc……..
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker.
Referred Documents :
Operation Standard "Annexure to ISW-QA-005"
page 1 of 2
Permissible Deviations for the Linear Height Gauge :
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (um)
1 | Length | Measured the permissible value on height checking gauge master. (Upward facing surface & Downward facing surface) |  | Calculation Method ±(1.1+0.6 L/600) µm
|  |  |  | Nominal |  | Acceptable
|  |  |  | 0 | - | ± 0
|  |  |  | 60 | ـ | ± 1.160µm
|  |  |  | 100 | ـ | ± 1.200µm
|  |  |  | 160 | ـ | ± 1.260µm
|  |  |  | 200 | ـ | ± 1.300µm
|  |  |  | 260 | ـ | ± 1.360µm
|  |  |  | 300 | ـ | ± 1.400µm
|  |  |  | 360 | ـ | ± 1.460µm
|  |  |  | 400 | ـ | ± 1.500µm
|  |  |  | 460 | ـ | ± 1.560µm
|  |  |  | 500 | ـ | ± 1.600µm
|  |  |  | 550 | ـ | ± 1.660µm
|  |  |  | 600 | ـ | ± 1.700µm
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | Mitutoyo | Linear Height Gauge
Checked By
Prepared By
Released By
Approved By
Checked By
Prepared By
Released By
Approved By

## Dial Calibration Tester
|  |  |  | Doc No : | ISW-QA-006
|  |  |  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.04.2015
Electronic Dial Calibration Tester
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For calibration of Dial Gauge ,Dial Test Indicator & Bore Gauge of various resolutions.
Reference Standards/Materials :
This document is based on maker's catalogue.
Description: | This machine could be used effectively for calibration of Dial Gauge,Dial Test Indicator & Bore Gauge of
different resolutions. This instrument consists of three parts viz Mechanical Stand , Probe P10 and
Electronic Indicating Unit D-80. The mechanical stand consists of robust cast iron base fixed with a
precision ground column with hard chrome plating. A measuring arm assembly can move up and down
smoothly on this column with the help of Rack and Pinion arrangement. The dial gauge is clamped with
the holder and probe is brought in contact with the Probe P 10. The electronic display is set to zero at
some specified value of dial graduation and then fixed nos of graduations are incremented and correspon-
-ding readings observed in the electronic display and compared with the nos of graduations advanced in
the dial gauge. The difference in value is compared with the standard specified in JIS and if lie within
specified range, Dial Gauge is accepted.
Standard Accessories :
The standard accessories required for the machine are Probe P 10, Electronic Display Unit D 80.
Calibration Frequency:
The annual calibration of the instrument is subject to satisfactory performance of the machine.
Calibration: | To ensure the performance, the instrument is calibrated by the External Calibration Agency.
Identification display method:
Calibration sticker on Dial Calibration Tester provided by Calibration Agency
Calibration Sticker provided by Calibration Agency
Recorded Document:
Calibration cerfificate given by External Calibration Agency after calibration.
Tracibility:
| NABL | EXTERNAL CALIBRATION AGENCY |  | DIAL CALIBRATION TESTER
Checked By
Prepared By
Approved By
Released By

## Outside Micrometer
[5 empty] | Doc No : | ISW-QA-007
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 6
[5 empty] | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Outside Micrometer
Calibration Method: |  |  | ISW-QA-007
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Micrometer Callipers for External Measurement.
2.0 References : |  | This document is based on JIS B 7502
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Slip Gauge Block of Grade 0 or Grade 1, Optical Flat Grade 1 or Grade 2
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 3
Permissible Deviations for Outside Micrometer :
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC :- | 0.001 mm
1 | Instrumental Error Verify at 2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80 & 25 mm combination of Gauge Blocks |  | Ensure the instrument has been made to read zero at the minimum measuring length.Next,insert a gauge block between the measuring faces,turn the ratchet and obtain the difference between the reading of Outside Micrometer and the dimension of the gauge block. |  | 0 < MR £ 25 | ± 2 microns
25 < MR £ 50
50 < MR £ 75
[5 empty] | 75 < MR £ 100 | ± 3 microns
100 < MR £ 125
125 < MR £ 150
[5 empty] | 150 < MR £ 175 | ± 4 microns
175 < MR £ 200
200 < MR £ 225
[5 empty] | 225 < MR £ 250 | ± 5 microns
2 | Flatness of Measuring Face |  | Allow an optical flat to contact closely with the measuring face & take the reading of number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. |  | 0 < MR £ 25 | 0.6 micron
25 < MR £ 50
50 < MR £ 75
75 < MR £ 100
100 < MR £ 125
125 < MR £ 150
150 < MR £ 175
175 < MR £ 200
200 < MR £ 225
225 < MR £ 250
3 | Parallelism of Measuring Face |  | Interpose a gauge block at four corners of the measuring face in turn, take readings of respective values by turning the rachet of the micrometer calipers & obtain the maximum difference thereof. |  | 0 < MR £ 25 | 2 microns
25 < MR £ 50
50 < MR £ 75
[5 empty] | 75 < MR £ 100 | 3 microns
100 < MR £ 125
125 < MR £ 150
150 < MR £ 175
[5 empty] | 175 < MR £ 200 | 4 microns
200 < MR £ 225
225 < MR £ 250
4 | Graduations |  | Visual |  | Good Contrast
5 | Thread Free Play |  | Visual |  | No perceptible free play (adjust threads if required)
6 | Constant Force Mechanism (if applicable) |  | Visual |  | Uniform Pressure
page 2 of 3
Identification display method:
Colour cable tie identification on Outside Micrometer as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Outside Micrometer in GCS.
Tracibility:
|  | NABL |  | EXTERNAL CALIBRATION AGENCY
Checked By
Prepared By
Approved By
Released By
OUTSIDE MICROMETER
MICROMETER CHECKER / GAUGE BLOCK
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Height Gauge
|  |  |  | Doc No : |  | ISW-QA-008
|  |  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : |  | 0
|  |  |  | Date : |  | 01.04.2015
Calibration of Height Gauge
Calibration Method: |  |  | ISW-QA-008
Calibration Frequency: |  |  | 1 month & 6 months
1.0 Scope: |  | This method describes the calibration of Height Gauges
2.0 References : |  | This document is based on JIS B 7517
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
ML:Measuring Length
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Caliper Checker , Gauge Blocks, Lever type Dial Indicator & Scriber.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 3
Permissible Deviations for the Height Gauge :
Permissible Deviations for the Height Gauge
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | L.C :- | 0.05 | 0.02 or 0.01
1 | Instrumental Error |  | Ensure that the Instrument has been made read zero at minimum measuring length. Next measure by putting the measuring surface of scriber of the height gauge on the gauge block of the caliper checker and the readings thus observed are substracted from the gauge block reading. |  | ML>=50 | ± 0.05 | ± 0.02
[5 empty] | 50< ML£100 | ± 0.06 | ± 0.03
[5 empty] | 100< ML£200 | ± 0.07
[5 empty] | 200< ML£300 | ± 0.08 | ± 0.04
[5 empty] | 300< ML£400 | ± 0.09
[5 empty] | 400< ML£500 | ± 0.10 | ± 0.05
[5 empty] | 500< ML£600 | ± 0.11
[5 empty] | 600< ML£700 | ± 0.12 | ± 0.06
[5 empty] | 700< ML£800 | ± 0.13
2 | Flatness of reference surface of base |  | Place the height gauge on two pieces of gauge blocks equal in nominal dimensions fixed on a rigid surface plate. Apply the measuring point of lever type dial gauge to the reference surface of base and move the lever type dial gauge through out the refering base and obtain the difference between the maximum value and minimum value obtained. |  | ---- | 0.005 mm | 0.005 mm
3 | Parallelism of reference surface of base to measuring surface of scriber |  | Place the height gauge on a precision surface plate, apply the measuring point of lever type dial gauge attached to the stand to the measuring surface of scriber and move the dial gauge covering the whole measuring surface of scriber and obtain the difference between the maximum value and the minimum value observed. |  | ---- | 0.010 mm | 0.010 mm
4 | Graduations (if applicable) |  | Visual |  | Good Contrast
Recorded Document:
Calibration Record of Height Gauge in GCS.
page 2 of 3
Identification display method:
Colour Tepra Tape identification of Height Gauge (1 month) as per method provided by MAP-ID.
Colour Paint Marker identification of Height Gauge (6 month) as per method provided by MAP-ID.
| Tepra Tape for 1 month frequency |  |  | Paint Marker for 6 month frequency
Tracibility:
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
DIAL TEST INDICATOR
CALIPER CHECKER, DIAL CALIBRATION TESTER & GAUGE BLOCKS
EXTERNAL CALIBRATION AGENCY
NABL
HEIGHT GAUGE
Checked By
Prepared By
Approved By
Released By

## Plunger Dial Gauge
[6 empty] | Doc No : | ISW-QA-009
Musashi Auto Parts India Pvt Ltd | [5 empty] | Issue No : | 6
[6 empty] | Rev No : | 0
[6 empty] | Date : | 01.04.2015
Calibration of Plunger Dial Gauge
Calibration Method: |  |  | ISW-QA-009
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Plunger Dials of LC 0.010 mm and 0.001 mm.
2.0 References : |  | This document is based on JIS B 7503
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Dial Calibration Tester with Probe & Electronic Indicating Unit D 80.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 of 3
Identification display method:
Colour Paint Marker identification on Plunger Dial Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Plunger Dial Gauge in GCS.
Tracibility:
|  | NABL |  | EXTERNAL CALIBRATION AGENCY
Checked By
Prepared By
Approved By
Released By
PLUNGER DIAL GAUGE
DIAL CALIBRATION TESTER
Checked By
Prepared By
Approved By
Released By

## Plunger Dial Gauge tole
Permissible Deviations for the Plunger Dial Gauge :
Sr No | Test Characteristics | Test Method | LC : | 0.010 mm | 0.002 mm |  | 0.001 mm
|  |  | MR (in mm) : | 10 mm Max | 2 mm Max | 2 < MR £ 10 | 1 mm Max | 1 < MR £ 2 | 2 < MR £ 5
1 | Indication Error | Mount Dial firmly on Dial Calibration Tester. Ensure movement axis in line with moving device. Set the reading of Dial Gauge at zero point. Now move the drum for 1/10 by 1/10 revolution upto two revolutions from zero point, 1/2 by 1/2 revolution upto five revolutions and 1 by 1 revolution upto the end point of the measuring range after exceeding five revolutions in upward direction and observe the readings from display unit, and returning back the plunger in downward direction, read the same measuring points as in upward direction and observe the readings. Now calculate the mentioned errors as shown in Indication Error diagram below. | Accuracy over an interval of any 1/10th revolution (Adjacent Error) | 8 microns | 4 microns | 5 microns | 2.5 microns | 4 microns | 5 microns
|  |  | Accuracy over an interval of any 1/2 revolution | ±9 microns | ± 5 microns | ± 6 microns | ± 3 microns | ± 5 microns | ± 6 microns
|  |  | Accuacy over an interval of any 1 revolution | ± 10 microns | ± 6 microns | ± 7 microns | ± 4 microns | ± 6 microns | ± 7 microns
|  |  | Accuracy over an interval of any 2 revolution | ± 15 microns | ± 6 microns | ± 8 microns | ± 4 microns | ± 6 microns | ± 8 microns
|  |  | Accuracy over whole measuring range | ± 15 microns | ± 7 microns | ± 12 microns | ± 5 microns | ± 7 microns | ±10 microns
2 | Retrace Error | Now find the maximum value of differences of the indications for the same measuring quantity when the plunger moves upwards and downwards. | ___ | 5 microns | 3 microns | 4 microns | 3 microns | 3 microns | 4 microns
3 | Repeatability | Sustain dial gauge so as to make the measuring probe parallel to the upper surface of measuring base, then move the measuring probe rapidly or slowly 5 times at any particular position in the measuring range, and find the largest difference b/w indicated values at respective times. | ___ | 5 microns | 0.5 microns | 1 microns | 0.5 microns | 0.5 microns | 1 microns
4 | Single Revolution Confirmation | Visual | Pointer should not take more than one revolution (For One Revolution Dial Only)
5 | General Feel of Dial |  | Smooth movement,No stickness,Clear Graduations
Indication Error Diagram :
page 2 of 3
Prepared By
Checked By
Approved By
Released By

## Vernier Caliper
[5 empty] | Doc No : |  | ISW-QA-010
[5 empty] | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : |  | 0
[5 empty] | Date : |  | 01.04.2015
Calibration of Vernier Caliper
Calibration Method: |  |  | ISW-QA-010
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Vernier calipers of resolution 0.10, 0.050, 0.020 &
0.010 mm with a maximum measuring range of 1000 mm.
2.0 References : |  | This document is based on JIS B 7507
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Enviorment :
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Caliper Checker.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for the Vernier Caliper :
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC:- |  | 0.10 or 0.05 | 0.020 or 0.010
1 | Instrumental Error for External Measurement & Internal Measurement |  | Ensure the instrument has been made to read zero at minimum measuring distance. Next check dimensions of various blocks of Caliper checker at three points of measuring jaws for External Measurement & Internal Measurement and obtain the difference between the dimension of block and reading of Vernier Caliper. |  | ML>=50 | ± 0.05 | ± 0.05 | ± 0.02
[5 empty] | 50 < ML £ 100 | ± 0.06 | ± 0.06 | ± 0.03
[5 empty] | 100 < ML £ 200 | ± 0.07 | ± 0.07
[5 empty] | 200 < ML £ 300 | ± 0.08 | ± 0.08 | ± 0.04
[5 empty] | 300 < ML £ 400 | ± 0.09 | ± 0.09
[5 empty] | 400 < ML £ 500 | ± 0.10 | ± 0.10 | ± 0.05
[5 empty] | 500 < ML £ 600 | ± 0.11 | ± 0.11
[5 empty] | 600 < ML £ 700 | ± 0.12 | ± 0.12 | ± 0.06
[5 empty] | 700 < ML £ 800 | ± 0.13 | ± 0.13
[5 empty] | 800 < ML £ 900 | ± 0.14 | ± 0.14 | ± 0.07
[5 empty] | 900 < ML £ 1000 | ± 0.15 | ± 0.15
Identification display method:
Colour Paint Marker on Vernier Caliper as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Varnier Caliper in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | CALIPER CKECKER
Checked By
Prepared By
Approved By
Released By
VERNIER CALIPER
Checked By
Prepared By
Approved By
Released By

## Bench Center
[5 empty] | Doc No : | ISW-QA-011
[5 empty] | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Bench Center
Calibration Method: |  | ISW-QA-011
Calibration Frequency: |  | 1 year
Calibration: |  | To ensure the performance, the instrument is calibrated by the External Agency.
1.0 Scope: |  | This method describes the calibration of Bench Center with Center Height 125 mm to
300 mm.
2.0 References : |  | This document is based on IS 5980-1978
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
Center Height (H): Height of Centers from the top face of the bed.
Center Distance: Distance between two centers.
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Test Mandrel (Straight), Magnetic Stand and Dial Test Indicator.
6.0 Calibration Process:
| 6.1 | The appearance ,cleaniness and condition of the centers(no dents, no damage and no
worn out) of the Bench Center are verified.
| 6.2 | The Push Handle of Bench Center should be working properly.
| 6.3 | The Head and Tailstock should be sliding smoothly.
| 6.4 | Bench Center is calibrated for the Coaxiality of Centers and parallelism of the Axis
of centers with respect to Guideways.
Prepared By |  | Checked By |  | Approved By |  | Released By
|  | Checked By |  |  | Approved By
page 1 of 3
Permissible Deviations for the Bench Center :
Sr No | Test Characterstic | Method |  | Acceptence Limit of Error (mm)
|  |  |  | Center Height:- |  | 125 and 160 (mm) | 200 to 300 (mm)
1 | Co-axiality of Centers | A Cylendrical test mandrel is held between the centers. A suitable dial indicator is mounted on the slide. The Dial Indicator is moved and the readings are taken in two planes perpendicular to each other close to the ends of the mandrel and also at intervals of 300 mm. The length of the mandrel shall be nearly equal to center distance. The difference in values is obtained and shall not exceed the permissible value. |  | Over any length of 300 mm between Centers |  | 0.01 | 0.015
|  |  |  | Center Distance upto 500 mm |  | 0.015 | 0.02
|  |  |  | Center Distance above 500 mm |  | 0.02 | 0.04
2 | Parallelism of the Axis of centers with respect to Guideways | A cylendrical test mandrel of 300 mm length with taper shank is inserted into the taper bore of the center. A suitable dial indicator is mounted on the base of bench center. Readings are taken in two planes perpendicular to each other at two positions close to the ends of the mandrel. The difference in values is obtained and shall not exceed the permissible value. This is repeated for the other center bore and at different positions along the length of base. |  | Per 300 mm leaning towards the free end of the mandrel |  | 0.01 | 0.015
3 | General Feel of Instrument |  |  | Smooth Movement , No Stickness,No Center Damage, No Dents etc…
Prepared By |  | Checked By |  | Approved By |  | Released By
|  | Checked By |  |  | Approved By
page 2 of 3
Calibration Certificate:
The following measuring items shall be contained in the calibration certificate.
a) | Co-axiality of Centers.
b) | Parallelism of the Axis of centers with respect to Guideways.
Identification display method:
Calibration sticker on Bench Center Provided by Calibration Agency
Calibration Sticker provided by Maker
Recorded Document:
Calibration cerfificate given by External Calibration Agency after calibration.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | BENCH CENTER
Prepared By |  | Checked By |  | Approved By |  | Released By
|  | Checked By |  |  | Approved By

## Bevel Protractor
[6 empty] | Doc No : |  | ISW-QA-012
[6 empty] | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : |  | 0
[6 empty] | Date : |  | 01.04.2015
Calibration of Bevel Protractor
Calibration Method: |  |  | ISW-QA-012
Calibration Frequency: |  |  | 1 year
1.0 Scope: |  | This method describes the calibration of Bevel Protractor
2.0 References : |  | This document is based on IS 4239-1970
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Coordinate Measuring Machine.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 of 2
Permissible Deviations for the Bevel Protractor:
Sr No | Measuring Item |  |  | Measuring Method |  |  | Acceptable limits of Error
1 | Anular Accuracies at 35°,95°,155°,220° and 360° |  |  | CMM |  |  | ± 5 minutes
Identification display method:
Colour Tepra Tape identification on Bevel Protractor as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Bevel Protector in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | CMM |  | BEVEL PROTECTOR
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Cylendrical Pins & Ball Anvils
|  |  |  | Doc No : | ISW-QA-013
|  |  |  | Issue No: | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.04.2015
Calibration of Cylendrical Pins, PCD Pins & Ball Anvils
Calibration Method: |  | ISW-QA-013
Calibration Frequency: |  | 1 year
1.0 Scope: |  | This method describes the calibration of Cylindrical Pins, PCD Pins & Ball Anvils
2.0 References : |  | This document is based on IS 11103-1984
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Indicating Micrometers (0.001 mm)
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Measuring Pins and Balls are calibrated for the pin & Ball diameter and observations are
recorded in the recording sheet in the respective boxes with respect to the size.
Permissible Deviations for the Cylindrical Pins, PCD Pins & Ball Anvils:
From 0.1 up to and including 20 mm dia = ± 2 microns
Traceability:
NABL |  | EXTERNAL CALIBRATION AGENCY | MICROMETER CHECKER | OUTSIDE MICROMETER |  | PINS & BALLS
Prepared By
Checked By
Approved By
Released By

## Taper Mandrel
|  |  |  | Doc No : | ISW-QA-014
|  |  |  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.04.2015
Calibration of Taper Mandrel
Calibration Method: |  |  | ISW-QA-014
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Taper Mandrels
2.0 References : |  | This document is based on Gauge Drawing.
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Bench Center and Plunger Dial of L.C 0.001 mm
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Mandrel measuring surface ,centers are cleaned and inspected for any damages,rust and
centers condition.
page 1 of 2
Permissible Deviations for the Test Bars/Taper Mandrel :
Sr No | Measuring Item |  | Measuring Method | Measuring Instrument | Acceptable limits of Error
1 | Outside Diameter (D1) |  | Ensure the instrument has been made to read zero at the minimum measuring length. Next, measure the Diameter(D1) of Mandrel at position shown in figure. | Outside Micrometer | +0.005
2 | Outside Diameter (d1) |  | Now measure the Diameter(d1) of Mandrel at position shown in figure. | Outside Micrometer | -0.005
3 | Runout |  | Load the Mandrel between centers of Bench Center. Plunger Dial is placed on the mandrel surface and pressure is applied. Now mandrel is rotated gently for 360° and maximum deviation on Dial is recorded. The runout is checked at both ends and at middle position and the maximum value of three is recorded in GCS System. | Bench Center & Dial Indicator (0.001 mm) | 0.002 mm (For New Recieved)
0.005 mm (For Periodic Calibration)
Identification display method:
Colour cable tie identification on Taper Mandrel as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Taper Mandrel in GCS.
Tracibility:
NABL
EXTERNAL CALIBRATION AGENCY
MICROMETER CHECKER/
DIAL CALIBRATION TESTER
OUTSIDE MICROMETER DIAL INDICATOR & BENCH CENTER
TAPER MANDREL
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Dial Test Indicator
[6 empty] | Doc No : |  | ISW-QA-015
[6 empty] | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : |  | 0
[6 empty] | Date : |  | 01.04.2015
Calibration of Dial Test Indicator
Calibration Method: |  |  | ISW-QA-015
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Dial Test Indicators of LC 0.010 m & 0.002 mm.
2.0 References : |  | This document is based on JIS B 7533
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: |  | 20°C
| 2) | Humidity : |  | No Excessive Humidity
| 3) | Air Quality : |  | N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Dial Calibration Tester with Probe & Electronic Indicating Unit D 80
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for the Dial Test Indicators :
Sr. No. | Measuring Item |  | Measuring Method |  | Acceptable limits of Error
[5 empty] | L.C :- 0.010 mm |  |  | L.C :- 0.002 mm
[5 empty] | 0.50 mm | 0.80 mm | 1.00 mm | 0.20 mm | 0.28 mm
1 | Wide Range Forward Accuracy |  | Set the contact point perpendicular to the measuring surface of the probe. Move the measuring probe by each 0.1 mm scale for TI with LC 0.01 mm and by each 0.02 mm scale for TI with LC 0.002 mm in the clockwise direction from basal point to terminal point of measuring range and determine the accuracy or error value from the error diagram(shown in next page) by substracting the reading of measuring device from the pointer reading. |  | 5 µm | 8 µm | 10 µm | 3 µm | 3 µm
2 | Neighboring Error |  | Determine the accuracy or error value from the error diagram by substracting the reading of measuring device from the pointer reading. |  | 5 µm |  |  | 2 µm
3 | Backward Error |  | After completion of forward accuracy measurement, return the measuring probe in anticlockwise direction starting from the same condition and meausre readings at same measuring points as in the measurement of wide range forward accracy and determine the error from the obtained forward and backward error diagrams. |  | 3 µm |  |  | 2 µm
4 | Repeatability |  | Sustain dial gauge so as to make the measuring probe parallel to the upper surface of measuring base, then move the measuring probe rapidly or slowly 5 times at any particular position in the measuring range, and find the largest difference b/w indicated values at respective times. |  | 3 µm |  |  | 1 µm
Identification display method:
Colour Pant Marker identification on Dial Test Indicator as per method provided by MAP-ID.
Colour Coding
page 2 of 3
Recorded Document:
Calibration Record of Dial Test Indicator in GCS.
Error Diagram:
Tracibility:
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
DIAL TEST INDICATOR
DIAL CALIBRATION TESTER
EXTERNAL CALIBRATION AGENCY
NABL

## Bore Gauge
[7 empty] | Doc No : |  | ISW-QA-016
[7 empty] | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd | [6 empty] | Rev No : |  | 0
[7 empty] | Date : |  | 01.04.2015
Calibration of Bore Dial Gauge
Calibration Method: |  |  | ISW-QA-016
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Bore Dial Gauge of different measuring ranges.
2.0 References : |  | This document is based on JIS B 7515
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: |  | 20°C
| 2) | Humidity : |  | No Excessive Humidity
| 3) | Air Quality : |  | N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Dial Calibration Tester with Probe & Electronic Indicating, Plunger Dial 0.010 mm & Ring Gauge.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 3
Permissible Deviations for the Bore Dial Gauges:
Sr. No. | Measuring Item |  | Measuring Method |  |  | Acceptable limits of Error
Measuring Range
[6 empty] | 10~18.5 | 18~35 | 35~ 60 | 50~100 | 100~160
1 | Wide Range Accuracy |  | Support the cylinder gauge in the dial calibration tester firmly and push the plunger of the gauge 0.10 mm by 0.10 mm starting from the origin to the final limit of effective measuring range, calculate the error by subtracting the reading of Dial Indicator and bore gauge and obtain the accuracy or error from the error diagram. |  |  | 10 µm | 10 µm | 10 µm | 10 µm | 10 µm
2 | Adjacent Error |  | After determining the wide range accuracy error next determine the accuracy or error value from the error diagram by substracting the reading of measuring device from the pointer reading. |  |  | 4 µm | 4 µm | 4 µm | 4 µm | 4 µm
3 | Repeated Accuracy |  | Insert the plunger of bore gauge into the ring gauge, swing the barrel to make three backward and forward motions in the plane including the axes of plunger and barrel, read the minimum indication during each forward or backward motion, and remove the bore gauge from the ring. Repeat this procedure three times, and determine the difference between maximum and minimum values among the nine readings taken. When making the mesurements, press the bore gauge tips at thr same places on the inside wall of ring gauge. |  |  | 3 µm | 3 µm | 3 µm | 3 µm | 3 µm
Identification display method:
Colour Cable Tie identification on Bore Dial Gauge as per method provided by MAP-ID.
Colour Coding
page 2 of 3
Recorded Document:
Calibration Record of Bore Dial Gauge in GCS.
Error Diagram:
Tracibility:
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
BORE DIAL GAUGE
DIAL CALIBRATION TESTER & RING GAUGE
EXTERNAL CALIBRATION AGENCY
NABL

## Spline Plug Gauge
[6 empty] | Doc No : | ISW-QA-017
[6 empty] | Issue No : | 6
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : | 0
[6 empty] | Date : | 01.04.2015
Calibration of Spline Plug Gauge
Calibration Method: |  |  | ISW-QA-017
Calibration Frequency: |  |  | 1 month & 3 months
1.0 Scope: |  | This method describes the calibration of Spline Plug Gauge.
2.0 References : |  | Approved Gauge Drawings
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Cylinder Pins and Outside Micrometer, V-Micrometer & CMM.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Spline Gauges are cleaned and inspected for any damages and dents.
| 6.7 | Ensure & verify the cylindrical pins taken are accurate to its size using micrometer.
page 1 of 2
Permissible Deviations for Spline Arbor :
Sr. No. | Measuring Item |  | Measuring Method |  |  | Measuring Instrument | Acceptable limits of Error
1 | Over Pin Diameter |  | Ensure that the instrument has been made to read zero at the minimum measuring length. Now take a set of cylendrical pin and place each cylendrical pin on opposite spline location, and check the diameter by placing measuring surface of micrometer over cylendrical pin and observe the value. Check the OPD at entry, middle & top positions in both X-X and Y-Y directions. |  |  | Outside Micrometer, Cylendrical Pins | As specified in Drawing
2 | Major Diameter |  | Check the no of splines. For even no of splines use Outside Micrometer and for odd no of splines use V-Micrometer. Ensure that the instrument has been made to read zero at the minimum measuring length. Next locate the Instrument on the major dia of the gauge and obtain the value. |  |  | Outside Micrometer or V-Micrometer | As specified in Drawing
3 | Minor Diameter |  | Place the Gauge on CMM, and obtain the value of minor dia. |  |  | CMM | As specified in Drawing
Identification display method:
Colour Tepra Tape identification of Spline Plug Gauge (1 month) as per method provided by MAP-ID.
Colour Cable Tie identification of Spline Plug Gauge (3 month) as per method provided by MAP-ID.
Tepra Tape for 1 month frequency
Cable Tie for 3 month frequency
Recorded Document:
Calibration Record of Spline Pluge Gauge in GCS.
Tracibility:
|  | EXTERNAL CALIBRATION AGENCY |  | MICROMETER CHECKER, GAUGE BLOCKS & CMM
Checked By
Prepared By
Approved By
Released By
NABL
OUTSIDE MICROMETER, CYLENDRICAL PIN & V-MICROMETER
SPLINE PLUG GAUGE
Checked By
Prepared By
Approved By
Released By

## Spline Ring Gauge
[5 empty] | Doc No : | ISW-QA-018
[5 empty] | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Spline Ring Gauge
Calibration Method: |  |  | ISW-QA-018
Calibration Frequency: |  |  | 1 month & 3 months
1.0 Scope: |  | This method describes the calibration of Spline Ring Gauge.
2.0 References : |  | Gauge Drawings
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: |  | 20°C
| 2) | Humidity : |  | No Excessive Humidity
| 3) | Air Quality : |  | N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Cylinderical Pins, Gauge Blocks & CMM.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Spline Gauges are cleaned and inspected for any damages and dents.
| 6.7 | Ensure & verify the cylindrical pins taken are accurate to its size using micrometer.
page 1 of 2
Permissible Deviations for Spline Arbor :
Sr. No. | Measuring Item |  | Measuring Method |  | Measuring Instrument | Acceptable limits of Error
1 | Between Pin Dimension |  | Take a set of cylendrical pin as specified in drawing and place each cylendrical pin on opposite spline, and check the BPD over the entire Spline Length of the gauge in both X-X and Y-Y directions by inserting the suitable gauge blocks.Gauge Block should be inserted between pins with optimal force(gauge block should qualify with sliding movement). |  | Gauge Blocks & Cylendrical Pins | As specified in Drawing
2 | Major Diameter |  | Place the Gauge on CMM, and obtain the value of minor dia. |  | CMM | As specified in Drawing
3 | Minor Diameter |  | Place the Gauge on CMM, and obtain the value of minor dia. |  | CMM | As specified in Drawing
Identification display method:
Colour Tepra Tape identification of Spline Ring Gauge (1 month) as per method provided by MAP-ID.
Colour Paint Marker identification of Spline Ring Gauge (3 month) as per method provided by MAP-ID.
Tepra Tape for 1 month frequency |  |  |  | Paint Marker for 3 month frequency
Recorded Document:
Calibration Record of Spline Ring Gauge in GCS.
Tracibility:
|  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCKS & CMM
Checked By
Prepared By
Approved By
Released By
NABL
CYLENDRICAL PIN
SPLINE RING GAUGE
Checked By
Prepared By
Approved By
Released By

## Disc Micrometer
[6 empty] | Doc No : | ISW-QA-019
Musashi Auto Parts India Pvt Ltd | [5 empty] | Issue No : | 6
[6 empty] | Rev No : | 0
[6 empty] | Date : | 01.04.2015
Calibration of Tooth Thickness Micrometer
Calibration Method: |  |  | ISW-QA-019
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Tooth Thickness Micrometer.
2.0 References : |  | This document is based on JIS B 7502
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Micrometer Checker, Slip Gauge Block Grade 0 or Grade 1 & Optical Flat Grade 1 or Grade 2.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for Tooth Thickness Micrometer :
Sr No | Test Characteristics |  | Test Method |  |  | Acceptable Limits of Error
[6 empty] | LC :- | 0.001 mm
1 | Instrumental Error Verify at 2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80 & 25 mm combination of Gauge Blocks |  | Ensure the instrument has been made to read zero at the minimum measuring length.Next,insert a gauge block between the measuring faces,turn the ratchet and obtain the difference between the reading of Tooth Thickness Micrometer and the dimension of the gauge block. |  |  | 0 < MR £ 25 | ± 4 microns
25 < MR £ 50
[6 empty] | 50 < MR £ 75 | ± 6 microns
75 < MR £ 100
[6 empty] | 100 < MR £ 125 | ± 7 microns
125 < MR £150
2 | Flatness of Measuring Face |  | Allow an optical flat to contact closely with the measuring face & take the reading of number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. |  |  | 0 < MR £25 | 1 micron
25 < MR £50
50 < MR £ 75
75 < MR £100
[6 empty] | 100 < MR £ 125 | 1.3 micron
125 < MR £150
3 | Parallelism of Measuring Face |  | Interpose a gauge block in the middle of both measuring faces, turn the ratchet and take its reading. Next, interpose the gauge block at four corners successively, take readings thereof, and obtain the maximum difference. |  |  | 0 < MR £25 | 4 microns
25 < MR £50
[6 empty] | 50 < MR £ 75 | 6 microns
75 < MR £100
[6 empty] | 100 < MR £ 125 | 7 microns
125 < MR £150
4 | Graduations |  | Visual |  |  | Good Contrast
5 | Thread Free Play |  | Visual |  |  | No perceptible free play (adjust threads if required)
6 | Constant Force Mechanism (if applicable) |  | Visual |  |  | Uniform Pressure
Identification display method:
Colour cable tie identification on Tooth Thickness Micrometer as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Tooth Thickness Micrometer in GCS.
Tracibility:
|  | NABL |  | EXTERNAL CALIBRATION AGENCY
TOOTH THICKNESS MICROMETER
MICROMETER CHECKER / GAUGE BLOCK
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Height Master Gauge
|  |  |  | Doc No : | ISW-QA-020
|  |  |  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.04.2015
Calibration of Height Master Gauge
Calibration Method: |  |  | ISW-QA-020
Calibration Frequency: |  |  | 6 months
1.0 Scope: |  | This method describes the calibration of Height Master Gauge
2.0 References : |  | Part Drawings
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Gauge Blocks (Grade 1) , Comparator Stand, Digimatic Dial Gauge.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | The height master size should be equal to the mean value of the spec. of part.
page 1 of 2
Permissible Deviations for the Height Gauge Block :
Sr No | Test Characteristics |  | Test Method | Measuring Instrument | Acceptable Limits of Error
1 | Height |  | Prepare the combination of Gauge Blocks equivalent to the mean size of the part. Now place the Gauge Block on comparator stand unde thhe probe of Dial Gauge. Set the value of Dial Gauge at zero. Next insert the Height Master under the Dial Gauge and note down the deviation from zero to observe the height of Height Master. | Gauge Block, Comparator Stand & Digimatic Dial Gauge | Mean Value of spec of part ±0.002
2 | Parallelism |  | Insert the Height Master under the Dial Gauge, take readings at four corners successively and obtain the maximum difference thereof. | Comparator Stand & Digimatic Dial Gauge | 0.002
3 | Size Itching |  | Visual | Mean value of the part should be itched on Height Master.
Identification display method:
Colour Paint Marker identification on Height Master Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Height Master Gauge in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY
Checked By
Prepared By
Approved By
Released By
DIAL GAUGE
GAUGE BLOCK, COMPARATOR STAND & DIAL CALIBRATION TESTER
Checked By
Prepared By
Approved By
Released By
HEIGHT MASTER

## Plain Plug Gauge
[5 empty] | Doc No : | ISW-QA-021
[5 empty] | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Plain Plug Gauge
Calibration Method: |  |  | ISW-QA-021
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Plain Plug Gauges.
2.0 References : |  | This document is based on JIS B 7420:1997
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Gauge Blocks (Grade 1) or better Specification , Pasameter.
6.0 Calibration Process:
| 6.1 | The Inspection measuring & test Equipment and Master(if any) is ensured at Standard
Temperature (20°±1°) for an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Based upon the Max Metal Limit & Min Metal Limit of the gauge, the maximum and
minimum permissible size for the Go and Not Go Gauge are evaluated.
| 6.6 | Gauge Blocks combinations are made as per the Max and Min permissible size
obtained.
page 1 0f 2
Permissible Deviations for the Plain Plug Gauge :
The accepatability of the gauge is subject to gauge Go & No Go observations falling in between the Max and
Min permissible size.
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument | Acceptable limits of Error
1 | Outside Diameter |  | Carry out the measurements using Pasa Meter in two directions perpendicular each other at the tip and root. Exclude the ranges 1 mm from both ends of gauge portion for the measurement. |  | Pasa Meter, Gauge Block (Grade1) or better specification | Outside Dia should lie in between max & min permissible size calcuated based upon max metal limit and min metal limit of the gauge
2 | Circularity |  | Measure the difference in diameter at XX & YY direction from the out side diameter 1/2, and take the max. of the values so obtained as the circularity of gauge. |  | Pasa Meter | As per the Standard JIS B 7420
3 | Cylindercity |  | Take the 1/2 of the difference between the max. value and the mini. value of the diameters obtained by the measurment of outside diameter as the cylindercity of gauge. |  | Pasa Meter | As per the Standard JIS B 7420
Identification display method:
Colour cable tie identification on Plain Plug Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Plain Plug Gauge in GCS.
Tracibility:
|  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCK
Checked By
Prepared By
Approved By
Released By
PASAMETER
PLAIN PLUG GAUGE
NABL
Checked By
Prepared By
Approved By
Released By

## Snap Gauge
[6 empty] | Doc No : | ISW-QA-022
[6 empty] | Issue No : | 6
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : | 0
[6 empty] | Date : | 01.04.2015
Calibration of Snap Gauge/Gap Gauge
Calibration Method: |  |  | ISW-QA-022
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Snap Gauges.
2.0 References : |  | This document is based on JIS B 7420:1997
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Gauge Blocks (Grade 1) or better Specification.
6.0 Calibration Process:
| 6.1 | The Inspection measuring & test Equipment and Master(if any) is ensured at Standard
Temperature (20°±1°) for an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Based upon the Max Metal Limit & Min Metal Limit of the gauge,the maximum and
minimum permissible size for the Go and Not Go Gauge are evaluated .
page 1 0f 2
Permissible Deviations for the Snap Gauge/Gap Gauge :
The accepatability of the gauge is subject to gauge go & No Go observations falling in between the Max and Min
permissible size.
Sr No | Measuring Item |  | Measuring Method |  |  | Measuring Instrument | Acceptable limits of Error
1 | Size |  | Take the maximum size of the gauge block through which the gap gauge just passes by its working load only when appropriate gauge blocks are inserted in the gap gauge as the size of the gap gauge. |  |  | Gauge Block (Grade1) or better specification | Size should lie in between max & min permissible size calcuated based upon max metal limit and min metal limit of the gauge
2 | Parallelism |  | Confirm that the gauge block equal to the size of the gauge part obtained using the method by means of gauge block plus the tolerance on parallelism does not enter between the gauging surfaces by the working load of the gap gauge in any direction. This size should not exceed the maximum permissible limit. |  |  | Gauge Block (Grade1) or better specification | As per the Standard JIS B 7420
Identification display method:
Colour cable tie identification on Snap Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Snap gauge in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCK
Checked By
Prepared By
Approved By
Released By
SNAP GAUGE
Checked By
Prepared By
Approved By
Released By

## Plain Ring Gauge
[5 empty] | Doc No : | ISW-QA-023
[5 empty] | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Plain Ring Gauge
Calibration Method: |  |  | ISW-QA-023
Calibration Frequency: |  |  | 1 month & 3 months
1.0 Scope: |  | This method describes the calibration of Plain Ring Gauge.
2.0 References : |  | This document is based on JIS B 7420:1997
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Bore Dial Gauge, Plunger Dial Gauge (0.001 mm).
6.0 Calibration Process:
| 6.1 | The Inspection measuring & test Equipment and Master(if any) is ensured at Standard
Temperature (20°±1°) for an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Based upon the Max Metal Limit & Min Metal Limit of the gauge,the maximum and
minimum permissible size for the Go and Not Go Gauge are evaluated .
| 6.6 | Gauge Blocks combinations are made as per the Max and Min permissible size obtained.
| 6.7 | Dial Bore Gauge is set to read zero on gauge block combinations thus made.
page 1 0f 2
Permissible Deviations for the Plain Ring Gauge :
The accepatability of the gauge is subject to gauge Go & No Go observations falling in between the Max and
Min permissible size.
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument | Acceptable limits of Error
1 | Inside Diameter |  | Measure the diameter in two directions perpendicular each other in the cross section at tip and root by using Dial Bore Gauge. Exclude the portions 1 mm from both ends for the measurement |  | Bore Dial Gauge | Inside Dia should lie in between max & min permissible size calcuated based upon max metl. limit and min metl. limit of the gauge.
2 | Circularity |  | Measure the difference in diameter at each cross section obtained from the measurement of inside diameter 1/2, and take the maximum of the values so obtained as the circularity of gauge. |  | Bore Dial Gauge | As per the Standard JIS B 7420
| Cylindercity |  | Measure the 1/2 of the difference between the maximum value & the minimum value of all the diameters obtained by measurment of inside diameter as the cylindercity of gauge. |  | Bore Dial Gauge
Identification display method:
Colour Tepra Tape identification on Plain Ring Gauge (1 month) as per method provided by MAP-ID.
Colour Paint Marker identification on Plain Ring Gauge (3 months) as per method provided by MAP-ID.
Tepra Tape for 1 month frequency
Paint Marker for 3 month frequency
Recorded Document:
Calibration Record of Plain Ring Gauge in GCS.
Tracibility:
Checked By
Prepared By
Approved By
Released By
NABL
EXTERNAL CALIBRATION AGENCY
DIAL CALIBRATION TESTER
BORE DIAL GAUGE & DIAL INDICATOR
PLAIN RING GAUGE
Checked By
Prepared By
Approved By
Released By

## Width Gauge
[6 empty] | Doc No : | ISW-QA-024
[6 empty] | Issue No : | 6
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : | 0
[6 empty] | Date : | 01.04.2015
Calibration of Plain Width Gauge
Calibration Method: |  |  | ISW-QA-024
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Width Gauge.
2.0 References : |  | This document is based on JIS B 7420:1997
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Digimatic Outside Micrometer (LC 0.001 mm).
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be thoroughly cleaned.
| 6.5 | Based upon the Max Metal Limit & Min Metal Limit of the gauge,the maximum and
minimum permissible size for the Go and Not Go Gauge are evaluated.
page 1 0f 2
Permissible Deviations for the Width Gauge :
The accepatability of the gauge is subject to gauge Go & No Go observations falling in between the Max and Min
permissible size.
Sr No | Measuring Item |  | Measuring Method |  |  | Measuring Instrument | Acceptable limits of Error
1 | Thickness |  | Ensure the instrument has been made to read zero at the minimum measuring length. Next, carry out the measurements using Outside Micrometer at the tip and root and obtain the value. Exclude the ranges 1 mm from both ends of gauge portion for the measurement. |  |  | Digimatic Outside Micrometer | Thickness should lie in between max & min permissible size calcuated based upon max metal limit and min metal limit of the gauge.
Identification display method:
Colour cable tie identification on Width Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Width Gauge in GCS.
Tracibility:
|  | EXTERNAL CALIBRATION AGENCY |  | MICROMETER CHECKER/ GAUGE BLOCK
Checked By
Prepared By
Approved By
Released By
OUTSIDE MICROMETER
PLAIN WIDTH GAUGE
NABL
Checked By
Prepared By
Approved By
Released By

## Feeler Gauge
[5 empty] | Doc No : | ISW-QA-025
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 6
[5 empty] | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Feeler Gauge
Calibration Method: |  |  | ISW-QA-025
Calibration Frequency: |  |  | 12 months
1.0 Scope: |  | This method describes the calibration of Feeler Gauge.
2.0 References : |  | This document is based on JIS B 7524:1992
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Digimatic Outside Micrometer (LC 0.001 mm)
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Feeler Gauges are checked for any damages.
| 6.7 | The feeler gauge blades surfaces and edges are checked for deformations,sharp edges
and smooth finish.
page 1of 2
Permissible Deviations for the Feeler Gauge :
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument | Acceptable limits of Error
1 | Thickness |  | Carry out the measurements using indicating micrometers at five points at equal intervals in the longitudinal directions excluding 20 mm approx from both ends. |  | Indicating Micrometer | Thickness Range | Tolerance on Thickness
[6 empty] | 0.01 ≤ TR £ 0.06 | ± 0.003
[6 empty] | 0.06 < TR £ 0.10 | ± 0.004
[6 empty] | 0.10 < TR £ 0.35 | ± 0.005
[6 empty] | 0.35 < TR £ 0.65 | ± 0.008
[6 empty] | 0.65 < TR £ 3.0 | ± 0.010
Identification display method:
Colour Tepra Tape identification of Feeler Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Feeler Gauge in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | MICROMETER CHECKER
Checked By
Prepared By
Approved By
Released By
OUTSIDE MICROMETER
Checked By
Prepared By
Approved By
Released By
FEELER GAUGE

## Try Square
[5 empty] | Doc No : |  | ISW-QA-026
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : |  | 6
[5 empty] | Rev No : |  | 0
[5 empty] | Date : |  | 01.04.2015
Calibration of Try Square
Calibration Method: |  |  | ISW-QA-026
Calibration Frequency: |  |  | 12 months
1.0 Scope: |  | This method describes the calibration of Try Square.
2.0 References : |  | This document is based on JIS B 7526:1995
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Coordinate Measuring Machine
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Try Squares are checked for any damages on reference faces.
| 6.7 | Try Square is calibrated for the squareness with reference to the base.
| 6.8 | Try Square is placed on the CMM and the squareness of outside and inside working
faces is checked by aligning the base.
page 1of 2
Permissible Deviations for the Try Square:
Sr No | Description of Try Square |  |  | MEASURING METHOD |  | Permissible deviation for Squareness (in microns)
Length * Height
1 | 100 * 70 |  |  | CMM |  | ± 3.0 microns
2 | 150 * 100 |  |  |  |  | ± 3.5 microns
3 | 200 * 130 |  |  |  |  | ± 4.0 microns
4 | 300 * 200 |  |  |  |  | ± 5.0 microns
Identification display method:
Colour Tepra Tape identification of Try Square as per method provided by MAP-ID.
Recorded Document:
Calibration Record Try Square in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | CMM
Checked By
Prepared By
Approved By
Released By
TRY SQUARE
Checked By
Prepared By
Approved By
Released By

## Steel Parallel
[5 empty] | Doc No : | ISW-QA-027
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 6
[5 empty] | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Steel Parallel
Calibration Method: |  |  | ISW-QA-027
Calibration Frequency: |  |  | 12 months
1.0 Scope: |  | This method describes the calibration of Steel Parallel.
2.0 References : |  | This document is based on JIS B 7514:1977
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Coordinate Measuring Machine,Indicating Micrometer
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Steel Parallel are checked for any damages on reference faces.
page 1 of 2
Permissible Deviations for the Steel Parallels:
Sr No | Measuring Item |  | Measuring Method | Meas. Instrument. | Acceptable limits of Error
[5 empty] | Effective Length (mm) | Tolerance
1 | Straightness of Working face |  | Steel Parallel is placed on the CMM and straightness of the working faces is checked by alligning the base. | CMM | 300 | 0.003
[5 empty] | 500 | 0.004
[5 empty] | 1000 | 0.006
[5 empty] | 1500 | 0.008
[5 empty] | 2000 | 0.01
[5 empty] | 3000 | 0.014
2 | Height Unevenness |  | Measure the height of the starightedge by using an indicating micrometer and obtain the difference between the largest and smallest values. | Indicating Micrometer | 300 | 0.003
[5 empty] | 500 | 0.004
[5 empty] | 1000 | 0.006
[5 empty] | 1500 | 0.008
[5 empty] | 2000 | 0.01
[5 empty] | 3000 | 0.014
3 | Parallelism of Side Face |  | Place the Steel Parallel on the precise surface plate and using the CMM/ Digital Height Gauge, determine the parallelism | CMM/Digital Height Gauge | 300 | 0.03
[5 empty] | 500 | 0.04
[5 empty] | 1000 | 0.06
[5 empty] | 1500 | 0.08
[5 empty] | 2000 | 0.1
[5 empty] | 3000 | 0.14
4 | Perpendicularity of Working face to Side Face |  | Employ a puppy dial gauge in the Digital Height Gauge. Place the Steel Parallel on the precise surface plate with its working face as a base and move the dial in the direction perpendicular to the working face keeping the dial pointer in contact with the side face. | CMM/Digital Height Gauge | 300 | 0.03
[5 empty] | 500 | 0.03
[5 empty] | 1000 | 0.04
[5 empty] | 1500 | 0.05
[5 empty] | 2000 | 0.06
[5 empty] | 3000 | 0.07
Identification display method:
Colour Tepra Tape identification on Steel Parallel as per method provided by MAP-ID.
Recorded Document:
Calibration Recorde of Steel Parallel in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY
VERNIER CHECKER/ DIAL CALIBRATION TESTER
O/S MICROMETER, HEIGHT GAUGE/ DAIL INDICATOR/ CMM
STEEL PARALLEL
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Magnetic V Block
[5 empty] | Doc No : |  | ISW-QA-028
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : |  | 6
[5 empty] | Rev No : |  | 0
[5 empty] | Date : |  | 01.04.2015
Calibration of Magnetic V-Block
Calibration Method: |  |  | ISW-QA-028
Calibration Frequency: |  |  | 12 months
1.0 Scope: |  | This method describes the calibration of Magnetic V Block.
2.0 References : |  | This document is based on JIS B 7540:1972
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Surface Plate Granite (Grade 1or higher), Digital Height Gauge,
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
faces is checked by aligning the base.
page 1 of 2
Permissible Deviations for the Magnetic V Block:
Permissible Deviations for the Magnetic V Block with designation under 100
Sr No | Measuring Item |  | Measuring Method |  | Meas. Instrument. |  | Acceptable Limits of Error
[7 empty] | Grade1 | Grade2
1 | Flatness of Bottom Surface |  | Place the upper surface of V Block on surface plate and bring in contact the dial pointer on the bottom surface to read zero at any one position.Then take the readings on the every part of bottom surface to determine the flatness of the bottom surface by the difference of readings. |  | Surface Plate Granite (Grade 1 or higher), Height Gauge, Dial Indicator (0.001 mm) |  | 0.010 mm | 0.020 mm
2 | Parallelism between bottom surface & cylinder on V surface |  | Place the bottom surface of V block on surface plate,support staight mandrel on each of V surface and determine the parallelism in length of V Block by the difference of readings of Dial Indicator. |  | Surface PlateGranite (Grade1or higher), Height Gauge, Dial Indicator (0.001 mm), Straight Mandrel |  | 0.010 mm | 0.020 mm
3 | Inclination of V Groove to bottom surface |  | Place the V Block on the Granite surface of CMM and allign the bottom surface of the block.Determine the angle between the inclined V surface with the bottom surface. |  | CMM |  | 45° approximately
4 | Parallelism between side surface and cylinder on V surface |  | Place the side surface of V block on surface plate,support staight mandrel on V surface in each and determine the parallelism in length of V Block from the difference of readings of Dial Indicator. |  | Surface PlateGranite (Grade1or higher), Height Gauge, Dial Indicator (0.001 mm), Straight Mandrel |  | 0.020 mm | 0.040 mm
Identification display method:
Colour Tepra Tape identification on Megnetic V-Block as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Recorde of Megnetic V-Block in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | VERNIER CHECKER/ DIAL CALIBRATION TESTER |  | HEIGHT GAUGE/ DAIL INDICATOR/ CMM
MAGNETIC V BLOCK
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Air Plug Gauge
[5 empty] | Doc No : | ISW-QA-029
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 6
[5 empty] | Rev No : | 0
[5 empty] | Date : | 01.04.2015
Calibration of Air Plug Gauge
Calibration Method: |  |  | 0
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Air Plug gauge
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
Balancing of Jets : The position of the Air Plug Gauge in the workpiece which reduces the air
flow in one jet must create a proportionate increase in air flow from an opposing jet.This
means that airflow and characteristics of the two jets must be "BALANCED"
Centrallity : Air Plug Gauge Jet diameters and nozzle drops must be on a common centerline
with the outer diameter of the plug gauge itself.
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Setting Masters, Read Out Unit, Bore Gauge & Slip Gauge Block Grade 0
6.0 Calibration Process:
| 6.1 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.2 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.3 | The test instrument should be throughly cleaned.
| 6.4 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
Prepared By |  |  | Checked By | Approved By |  | Released By
page 1 of 3
Permissible Deviations for the Air Plug Gauge:
Sr. No. | Measuring Item |  | Measuring Method |  | Meas. Instrument. | Acceptable Limits of Error
1 | Setting Master Diameter |  | Set the Bore Gauge at zero position on Bore Gauge Setting Block (Gauge Block combination equivalent to the Dia as of the setting ring). Next insert the Bore Gauge inside the setting ring and swing the barrel in forwrd and backward direction and observe the minimum value as the diameter of the setting ring. |  | Bore Gauge, Bore Gauge Setting Block & Slip Gauge Block Grade 0 | ±0.002 mm
2 | Nozzle Balancing Checking |  | 1) Calibrate the air gauging system using the max and min masters in a normal manner. 2) Orient the gauge with a horizontal centerline so the jets are situated vertical and located midway within the width of maximum setting master. Note the exact reading. 3) Reposition the setting master so that the nozzles are reversed 180 degrees. Observe the difference between the indicator reading from step 2 above. |  | Air Plug Gauge, Setting Masters, Read out Unit (Digital,Dial Type) | Differnece up to max one division on the scale (max 1 micron) for new bought out parts. This difference could be allowed till max 3 divisions in periodic checking.
3 | Centrality Checking |  | 1) Calibrate the air gauging system using the max and min masters in a normal manner. 2) Orient the gauge with a horizontal centerline so the jets are situated horizontal and located midway within the width of maximum setting master.Note the exact reading. 3) Rotate the setting master so that the jets are reversed 180 degrees. Observe the difference between the indicator reading from step 2 above. |  | Air Plug Gauge, Setting Masters, Read out Unit (Digital,Dial Type) | Differnece up to max two division on the scale (max 2 microns) for new bought out parts. This difference could be allowed till max 3 divisions during periodic calibration & inspection.
Prepared By |  | Checked By |  | Approved By |  | Released By
page 2 of 3
Identification display method:
Colour Paint Marker identification on Air Plug Gauge as per method provided by MAP-ID.
Colour Paint Marker identification on Setting Master as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Air Plug Gauge in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  | Released By
EXTERNAL CALIBRATION AGENCY
NABL
GAUGE BLOCK/ DIAL CALIBRATION TESTER
BORE GAUGE/ READ OUT UNIT
AIR PLUG GAUGE

## Cam Tester (ACCRETECH)
|  | Doc No : |  | ISW-QA-030
|  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  | Rev No : |  | 0
|  | Date : |  | 01.04.2015
Cam Tester (Accretech)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid
equipment, operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of Cam Profile of Various Cam Components..
Reference Standards/Materials :
Master Cam (MF1001-3-F008), Height Master Gauge(MF1001-4-G007), Gauge Blocks and
Electronic Dial Unit
Description: | This Machine could be used effectively for checking of various cam profile in both manual &
automatic mode. It is possible to program the machine for a specific component for different
values of lift angle and lift error as specified and could be saved with a distinguished identification
No. The whole software program is menu driven thereby making the machine convenient for use.
The detailed descriptions of the various menus can be referred from Operating Manual.
Standard Accessories:
The standard accessries required for the operation of the machine are Height Master Gauge,
Gauge Block , Cylindrical Pins, Electronic Dial Unit.
Calibration: | The annual calibration of the instrument is subject to satisfactory performance of the machine.
The master Cam piece are checked for Cam Profile on Periodic Basis and the results are
compared with the Master Graph taken during the Machine Installation time.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the instrument is calibrated by the maker .
Referred Document :
Operation Standard "Annexure to ISW-QA-030"
page 1 of 2
Permissible Deviations for the Cam Tester (Acctretech) :
Sr No | Test Characteristics |  | Acceptable Limits of Error (mm)
1 | Measurement accuracy (lift detector accuracy) |  | ±0.0015 mm
2 | Deviation of head stock center hole. |  | ±0.002 mm
3 | Parallelism of center hole with tabel and detum groove (with in vertical & horizontal both surface) |  | ±0.010/100mm
4 | Parallelism of contact mounting surface with both centers. |  | ±0.002/20mm
5 | Straightness of contactor Mounting surface ti table. |  | ±0.002/20mm
6 | Parallelism of lift detector (forward and backward movement) with table. |  | ±0.010/45mm
7 | Straightness of lift detector forward and backward movement (with in the verical & horizontal both surface) |  | ±0.004/45mm
8 | Parallelism of lift detector with both centers right to left movement. (with in the vertical & horizontal surface) |  | Vertical:- 0.010/100mm Horizontal:- 0.015/150mm
9 | Repetitive master cam Accuracy 5 times |  | 0.001 mm/5 times
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
Checked By
Prepared By
Approved By
Released By
NABL
ACCRETECH
CAM TESTER
Checked By
Prepared By
Approved By
Released By

## CNC Gear Checker
|  |  |  | Doc No : | ISW-QA-031
|  |  |  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.04.2015
CNC Gear Checker (CLP-35)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid
equipment,operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of Gear Profile/Lead for all types of spur/helical Gears and Shafts in the capacity
of max.outside Dia 350 mm with centers and max.shaft length 30~ 400 mm.
Reference Standards/Materials:
Master Involute & Lead
Test Mandrel (Straight)
Description: | Machine can be programmed for individual components with the base data availability of
Module, Pressure Angle,No of Teeth,Helix Angle,Gear Widthetc..The components are hold in
between the machine centers with clamping pressure automatically controlled by the
machine.Component is located between the centers using Dog Carrier.The Measuring Stylus
of the Detector Unit is brought in contact with the Gear Flank and program is run in automatic
mode.
The Lead & Profile Error thus evaluated will be printed once the measurement is over.It is also
possible to check the pitch error of the External & Internal Gear tooth.Unlike Manual gear
tester, the values of Lead & Profile Error is automatically calculated ending the measurement
ambiguity subjected with Manual Gear Tester.
Standard Accessories :
The standard accessries required for the operation of the machine are Stylus, Lower Center,
Dog and carrier,.
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done
using a Test Mandrel and Master Involute & Lead supplied along with the machine by the
maker .The master Involute and Lead is loaded between centers and lead and profile for
different helix angles installation.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-031"
Periodic Calibration of CNC Gear Checker (CLP-35)- "ISW-QA-031A"
page 1 of 2
Permissible Deviations for the CNC Gear Tester (CLP-35) :
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (mm)
1 | Runout of Main Shaft Reference Face |  |  | a. Radial Within; 0.003
b. Thrust Within; 0.003
2 | Lower Center Runout & Upper Center Runout |  |  | a. Lower Center Within; 0.003
b. Thrust Within; 0.003
3 | Concentricity and parallelism of both Centers. |  |  | 1) Concentricity 300 L & 150 L Within: 0.006
2) Parallelism 150 L Within: 0.006
4 | Parallelism of Detector Slide to Center Line of Both Center |  |  | 1) Parallelism X-Axis & Y-Axis Within 0.006
1) Straightness X-Axis & Y-Axis Within 0.003
5 | Master Measurment:- 1) Tengent Master 2) Master Gear |  |  | Compared with Master Value 1) FQ Within:- ± 0.003 2) FF Within:- ± 0.005
Variability Within:- 0.003
Tracibility:
| NABL | OSAKA SEIMITSU | CNC Gear Tester
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Coordinate Meas Machine
|  |  | Doc No : |  | ISW-QA-032
|  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : |  | 0
|  |  | Date : |  | 01.04.2015
Coordinate Measuring Machine (Crysta 574)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of all Complex Three Dimensional Measurements for all Critical Inspections Loactions which
is either not possible or very difficult to check with normal Routined Measuring Equipments.
Reference Standards/Materials:
1) Master Ball (19.9921 mm)
2) Judgment doc. No=0124501
Description: | Machine makes use of Geopack Software for basic evaluation & measurements of various geometrical
parameters. Machines could be programmed for individual components Inspection Locations with the
availability of Base Drawing as a reference.It is possible to check various geometrical dimensions and
analysis of the same can be done as desired & needed for Inspection.
Machine has a facility for both touch trigger probing and Scanning .
Standard Accessories :
The standard accessries required for the operation of the machine are Ruby Stylus
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.In case of
any doubt in the overall performance of the machine,the periodic calibration is done using a Straight Test
Bar. The Test Bar is subjected to Repeat checking of Diameter 25 Nos of Times and checked for
repeatability.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-032"
page 1 of 2
Permissible Deviations for the Coordinate Measuring Machine :
Sr No | Test Characteristics | Test Method | Acceptable Limits of Errors (µm)
1 | Error in direction E1 (x) | Measuring the direction of different axis use Length checking gauge master in different axis and measured the Errors in different length. | 30 | 150 | 250 | 170 | 490
Acceptable limits of error:1.9+4.0 L/1000
| Error in direction E2 (y) |  | 30 | 190 | 350 | 530 | 710
Acceptable limits of error:1.9+4.0 L/1000
| Error in direction E3 (z) |  | 30 | 110 | 210 | 290 | 390
Acceptable limits of error:1.9+4.0 L/1000
| Error in direction E4 |  | 30 | 110 | 310 | 510 | 710
Acceptable limits of error:1.9+4.0 L/1000
| Error in direction E5 |  | 30 | 110 | 310 | 510 | 710
Acceptable limits of error:1.9+4.0 L/1000
| Error in direction E6 |  | 30 | 110 | 310 | 510 | 710
Acceptable limits of error:1.9+4.0 L/1000
| Error in direction E7 |  | 30 | 110 | 310 | 510 | 710
Acceptable limits of error:1.9+4.0 L/1000
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
NABL
MITUTOYO
COORDINATE MEASURING MACHINE

## Roundness Tester
|  |  | Doc No : |  | ISW-QA-033
|  |  | Issue No : | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : | Rev No : | 0
|  |  | Date : | Date : | 01.04.2015
Roundness Tester (RA 2000DH)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of roundness and cylindricity of Round Gears and Shafts.
Reference Standards/Materials :
1) Origin Point Gauge
2) Reference Hemisphere
3) Optical Flat
4) Gauge Block Set
5) Judgment doc. No=MSA/10/RA/610
Description: | This machine requires alligning the workpiece such as centering,levelling with respect to the using
analysis item centering ,levelling or centering/leveling.Once workpiece is alligned with respect to the
worktable ,the analysis item such as Roundness or cylindricity is activated by double clicking the
respective icon and as per the instructions avaialble,the roundness and cylindricity values are
calculated. For detailed description ,the Roundness Tester manual can be referred.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus,
(Part No 12AAB681),Form Measuring Software supplied with the machine.
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done.
There are three types of calibration :
1) Static Calibration
2) Dynamic Calibration
3) Diameter measurement calibration
In case of ststic calibration ,a known displacement is given to the detector using a optical flat and
gauge blocks and the detector is calibrated to the specified value and the same is registered.
The Dynamic calibration and Diameter calibration are usually done by the maker in case of doubted
performance with the specified masters.
Origin Point Gauge is used for X and Z axis origin gauge measurement.For detailed description of
the same ,refer to the Roundness Tester manual.
Calibration Frequecy:
Once in a 3 Year & In case of suspected performance ,the machine is calibrated by the maker.
Referred Documents:
Operation Standard "Annexure to ISW-QA-033"
Periodic Calibration of Roundness Tester (RA-2000DH) - "ISW-QA-033A"
page 1 of 2
Permissible Deviations for the Roundness Tester :
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (µm)
1 | R-Axis Straightness | Measuring by the resting the Optical Flat on rotatry table and measure the R-Axis Straightness on Optical Flat on top side. |  | Straightness of R-Axis Should be With in 1µm / 150mm
3 | Z-Axis parallelism with tabel | Measuring by the resting the Cylinderical master on rotatry table and measure the Z-Axis parallelism on cylinderical master. |  | Parallelism of Z-Axis Should be With in 1µm / 300mm
4 | Roundness Measurement | Measuring by the resting the Hemispere on rotatery table & measure the roundness of Hemispere. |  | (0.02+5H/10000) <0.1µm
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | MITUTOYO | ROUNDNESS TESTER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Contour Tester
[5 empty] | Doc No : |  | ISW-QA-034
[5 empty] | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : |  | 0
[5 empty] | Date : |  | 01.04.2015
Contour Tester (CV-3000 S4)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of dimensional measurements such as angle,radius and distance measureemnts
of various machined components.
Reference Standards/Materials :
1) Standard Optical flat
2) X-Axis Inspection Standard
3) Gauge block Set
2) Judgment doc. No=218-487E
Description: | This machine makes use of Formpack software for measurement and evaluation of various
dimensional parameters viz angle,radius,linear dimension etc .The components are mounted in
the rotary vice and are positioned as per the requirement.The measuring stylus is made to travel
on the inspection location on the components and the contour is generated using the software. The
contour is then subjected to various analysis buttons and results are obtained.
Standard Accessories :
The standard accessories required for the operation of machine are Standard Stylus(Part No 354884),
Formpack Software,Cross Travel Table,Rotary Vise,XY Levelling Table with swivel etc….
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done.
There are three types of calibration :
1) Performing Z Axis gain adjustment.
- Static Calibration
- Dynamic Calibration
2) Performing straightness compensation measurement
3) Performing symmetric compensation measurement
4) Performing stylus radius measurement
In case of static calibration ,a levelling is performed for the calibration gauge,then one point on the
upper face and one point on the lower face of the gauge are measured.The Z axis gain adjustment
ratio is determined by calculating the step between the two points.
The stylus radius measurement is done using a pin gauge diamater.
The other calibration methods specified above are usually done by the maker in case of doubted
performance with the specified masters.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-034"
Periodic Calibration of Contour Tester (CV-3000 S4)- "ISW-QA-034A"
page 1 of 2
Permissible Deviations for the Contour Tester :
Sr No | Test Characteristics |  | Test Method |  |  | Acceptable Limits of Error (µm)
1 | Straightness of Drive unit |  | Measured Straightness Error of Drive unit on Optical Flat |  |  | Straightness Should be with in 2.0µm/200MM
2 | X-Axis Measuring Accuracy |  | Measured X-Axis Accuracy in 3 traverse length (25mm,50mm,75mm). |  |  | X-Axis measuring Accuracy should be with in ±(0.8+0.02L)µm
3 | Z-Axis measuring Accuracy |  | Measured Z-Axis Accuracy on Gauge block (40mm Step) |  |  | Z-Axis measuring Accuracy should be with in ±(1.6+¦2H¦/100)µm
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | MITUTOYO |  | CONTOUR TESTER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Roughness Tester Kosaka
|  |  |  | Doc No : |  | ISW-QA-035
|  |  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : |  | 0
|  |  |  | Date : |  | 01.04.2015
Kosaka Roughness Tester (SE-3500)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid
equipment, operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of various roughness parameters & Bibri of various machines components
Reference Standards/Materials :
1) SS-G21 Surface Finish Master S94
2) Judgment doc. No=D513007
Description: | This machine makes use of Stylus that traces waviness on a measuring surface.
The stylus is positioned on the component perpendicular to the machining direction parallel to
the detector and respective roughness parameters are evaluated (Part No PUDJ2S-60),Sine Vise.
Standard Accessories :
The standard accessories required for the operation of the machine are Standard Stylus,
(PUDJ2S-60)
Calibration: | The periodic calibration is done using Precision Reference Specimen described above.The
reference specimen is placed on the V-Block.The stylus is made to traverse a decided cut off
length and roughness value obserbed is compared with the master value specified on the
precision specimen.In case of any mismatch the conversion factor is applied and the machine is
recalibrated using the master specimen.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-0035"
Periodic Calibration of Roughness Tester (SJ-402) - "ISW-QA-0035"
page 1 of 2
Permissible Deviations for the Roughness Tester :
Sr No | Test Characteristics | Test Method |  |  | Acceptable Limits of Error
1 | Test side E on Specimen | Measuring the Specimen first adjust parallelism between the drive unit and the specimen and measure the surface finish. |  |  | Measuring value Should be with in ±5% of the indicated value
|  | Ra | ـ | 5µm
|  | Sm | ـ | 0.185mm
|  | P-P | ـ | 11.5µm
2 | Test Side C on Specimen | Measuring the Specimen first adjust parallelism between the drive unit and the specimen and measure the surface finish. |  |  | Measuring value Should be with in ±5% of the indicated value
|  | P-P | ـ | 0.47
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | KOSAKA LABORATORY |  | KOSAKA ROUGHNESS TESTER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## profile projector
|  |  | Doc No : |  | ISW-QA-036
|  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : |  | 0
|  |  | Date : |  | 01.04.2015
Profile Projector (V 12B)
Purpose: | This documented text aims to demonstrate the various features available in the above said equipment,
operating procedures, calibration methods and reference standards used ,if any.
Scope: | For checking of various linear 1D & 2 D dimensional parameters viz circle diameter, distance, angle
& chamfers etc…
Reference Standards/Materials :
1) Standard Slip
Description: | This machine could be used effectively for measurement of one dimensional and two dimensional
linear dimensions with great accuracy and repeatability. This measuring instrument could be used
to measure linear dimensions in the capacity of 150 mm Max length in X Axis & 100 mm Max
tissue length in Y Axis Before measurement., the instrument needs to clean properly with the help
of paper & Component should clamp properly on V-Block.
Standard Accessories :
The standard accessories required for the operation of the machine are Standard Lance, bulb,
Standard Slip etc……..
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine. In case
of any doubt in overall performance of machine periodic calibration is done using a standard slip. By
putting the standard slip on bed & compare the length of slip with machine output.
Calibration Frequency:
Once in a 2 Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-036"
page 1 of 2
Permissible Deviations for the Profile Projector :
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (µm)
1 | Distance | Measuring by the resting the std. slip on the glass and measure the distance of std. slip X-axis and Y-axis. (Use Different range std. slip) |  | Measuring value Should be with in ±3.2 µm of the indicated value.
2 | Angle | Measuring by the resting the angle master on the glass and measure the angle. |  | Measuring value Should be with in 6.1 Min of the indicated value.
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | A.V INDUSTRIES | PROFILE PROJECTOR
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Cam Tester (Adcole)
|  |  |  | Doc No : | ISW-QA-037
|  |  |  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 01.04.2015
Cam Tester (Adcole)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of Cam Profile of Various Cam Components..
Reference Standards/Materials :
Master Cam (C07L44GW011-005), Gauge Pins (P/N:TL05642)
Description: | This Machine could be used effectively for checking of various cam profile automatic mode.It is
possible to program the machine for a specific component for different values of lift angle and lift error
as specified and could be saved with a distinguished identification No.The whole software program is
menu driven thereby making the machine convenient for use.The detailed descriptions of the various
menus can be referred from Operating Automatic.
Standard Accessories:
The standard accessries required for the operation of the machine are Gauge Pins (P/N:TL05642).
Calibration: | The annual calibration of the instrument is subject to satisfactory performance of the machine.
The master Cam piece are checked for Cam Profile on Periodic Basis and the results are compared
with the Master Graph taken during the Machine Installation time.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Document :
Operation Standard "Annexure to ISW-QA-037"
page 1 of 2
Permissible Deviations for the Cam Tester (Adcole):
Sr No | Test Characteristics |  |  | Acceptable Limits of Error (inches)
1 | Diametric accuracy |  |  | Measured Diametric Accuracy shoulde be with in ±0.00020 inches maximum
2 | Spindle Run-Out |  |  | Measured Spindle run-out shoulde be with in 0.00015 inches maximum
3 | Angle Accuracy | 1.07 arcsec peak to peak |  | 10 arcsec peak to peak maximum
|  | 0.49 arcsec peak to peak |  | 1 arcsec peak to peak maximum
4 | Axial (Z Axis) Accuracy (Strain Gage Models Only) |  |  | 0.000100 inches per inch of travel maximum
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL |  | ADCOLE | CAM TESTER (ADCOLE)
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Contour Tester (CV-3100)
|  |  | Doc No : |  | ISW-QA-038
|  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : |  | 0
|  |  | Date : |  | 01.04.2015
Contour Tester (CV-3100 )
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of dimensional measurements such as angle,radius and distance measureemnts
of various machined components.
Reference Standards/Materials :
1) Std. Optical Flat
2) X-Axis Inspection Standard
3) Gauge Block Set.
2) Judgment doc. No=218-487E
Description : | This machine makes use of Formpack software for measurement and evaluation of various
dimensional parameters viz angle,radius,linear dimension etc .The components are mounted in
the rotary vice and are positioned as per the requirement.The measuring stylus is made to travel
on the inspection location on the components and the contor is generated using the software.The
contour is then subjected to various analysis buttons and results are obtained.
Standard Accessories :
The standard accessories required for the operation of the machine are Standard Stylus, (Part No
354884), Formpack Software,Cross Travel Table,Rotary Vise,XY Levelling Table with swivel,
Pin Gauge, Master Ball, Gauge Block Master etc….
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done.
| Batch calibration : | 1) Performing Z Axis gain adjustment
2) Performing straightness compensation measurement
3)Performing stylus radius measurement
In case of Batch calibration measurment complete then Registering the compensation result .
The other calibration methods specified above are usually done by the maker in case of doubted
performance with the specified masters.
Calibration Frequecy :
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-038"
Periodic Calibration of Contour Tester (CV-3100 S4)- "ISW-QA-038A"
page 1 of 2
Permissible Deviations for the Contour Tester (CV-3100):
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (µm)
1 | Straightness of Drive unit | Measured Straightness Error of Drive unit on Optical Flat |  | Straightness Should be with in 2.0µm/200MM
2 | X-Axis Measuring Accuracy | Measured X-Axis Accuracy in 3 traverse length (25mm,50mm,75mm). |  | X-Axis measuring Accuracy should be with in ±(0.8+0.02L)µm
3 | Z-Axis measuring Accuracy | Measured Z-Axis Accuracy on Gauge block (40mm Step) |  | Z-Axis measuring Accuracy should be with in ±(1.6+¦2H¦/100)µm
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | MITUTOYO | CONTOUR TESTER (CV-3100)
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Contour Tester (CV-3200)
|  |  | Doc No : |  | ISW-QA-039
|  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : |  | 0
|  |  | Date : |  | 01.04.2015
Contour Tester (CV-3200 )
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of dimensional measurements such as angle,radius and distance measurements
of various machined components.
Reference Standards/Materials :
1) Std. Optical Flat
2) X-Axis Inspection Standard
3) Gauge Block Set
2) Judgment doc. No=218-487E
Description : | This machine makes use of Formpack software for measurement and evaluation of various
dimensional parameters viz angle,radius,linear dimension etc .The components are mounted in
the rotary vice and are positioned as per the requirement.The measuring stylus is made to travel
on the inspection location on the components and the contor is generated using the software.The
contour is then subjected to various analysis buttons and results are obtained.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus, (Part No
354884),Formpack Software,Cross Travel Table,Rotary Vise,XY Levelling Table with swivel,
Pin Gauge, Master Ball, Gauge Block Master etc….
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done.
| Batch calibration : | 1) Performing Z Axis gain adjustment
2) Performing straightness compensation measurement
3) Performing stylus radius measurement
In case of Betch calibration measurment complet then Registering the compensation result .
The other calibration methods specified above are usually done by the maker in case of doubted
performance with the specified masters.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents :
Operation Standard "Annexure to ISW-QA-039"
Periodic Calibration of Contour Tester (CV-3100 S4)- "ISW-QA-039A"
page 1 of 2
Permissible Deviations for the Contour Tester (CV-3100):
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (µm)
1 | Straightness of Drive unit | Measured Straightness Error of Drive unit on Optical Flat |  | Straightness Should be with in 2.0µm/200MM
2 | X-Axis Measuring Accuracy | Measured X-Axis Accuracy in 3 traverse length (25mm,50mm,75mm). |  | X-Axis measuring Accuracy should be with in ±(0.8+0.02L)µm
3 | Z-Axis measuring Accuracy | Measured Z-Axis Accuracy on Gauge block (40mm Step) |  | Z-Axis measuring Accuracy should be with in ±(1.6+¦2H¦/100)µm
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | Mitutoyo | Contour Tester (CV-3200)
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Roundness Tester(RA-2200)
|  |  |  | Doc No : |  | ISW-QA-040
|  |  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : |  | 0
|  |  |  | Date : |  | 01.04.2015
Roundness Tester (RA 2200)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of roundness and cylindricity of Round Gears and Shafts
Reference Standards/Materials :
1) Origin Point Gauge
2) Reference Hemisphere
3) Optical Flat
4) Gauge Block Set
5) Judgment doc. No=MSA/10/RA/610
Description: | This machine requires alligning the workpiece such as centering,levelling with respect to the using
analysis item centering ,levelling or centering/leveling.Once workpiece is alligned with respect to the
worktable , the analysis item such as circularity or cylindricity is activated by double clicking the
respective icon and as per the instructions avaialble, the roundness and cylindricity values are
calculated.
For detailed description ,the Roundness Tester automatic can be referred.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus, (Part No
12AAB681),Form Measuring Software supplied with the machine.
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done.
| There are three types of calibration : |  | 1) Static Calibration
2) Dynamic Calibration
3) Diameter measurement calibration
In case of ststic calibration ,a known displacement is given to the detector using a optical flat and
gauge blocks and the detector is calibrated to the specified value and the same is registered.
The Dynamic calibration and Diameter calibration are usually done by the maker in case of doubted
performance with the specified masters.
Origin Point Gauge is used for X and Z axis origin gauge measurement.For detailed description of
the same ,refer to the Roundness Tester manual.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents:
Operation Standard "Annexure to ISW-QA-040"
Periodic Calibration of Roundness Tester (RA-2000DH) - "ISW-QA-040A"
page 1 of 2
Permissible Deviations for the Roundness Tester :
Sr No | Test Characteristics | Test Method |  |  | Acceptable Limits of Error (mm)
1 | R-Axis Straightness | Measuring by the resting the Optical Flat on rotatry table and measure the R-Axis Straightness on Optical Flat on top side. |  |  | Straightness of R-Axis Should be With in 1µm / 150mm
2 | Z-Axis parallelism with tabel | Measuring by the resting the Cylinderical master on rotatry table and measure the Z-Axis parallelism on cylinderical master. |  |  | Parallelism of Z-Axis Should be With in 1µm / 300mm
3 | Roundness Measurement | Measuring by the resting the Hemispere on rotatery table & measure the roundness of Hemispere. |  |  | (0.02+5H/10000) <0.1µm
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | Mitutoyo |  | Roundness Tester
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Roundness Tester(Kosaka)
|  |  |  | Doc No : |  | ISW-QA-041
|  |  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : |  | 0
|  |  |  | Date : |  | 01.04.2015
Roundness Tester (Kosaka)
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of roundness and cylindricity of Round Gears and Shafts
Reference Standards/Materials :
1) Point Gauge Master
2) Reference Hemisphere
3) Standard magnification Speciman
Description: | This machine requires alligning the workpiece such as centering,levelling with respect to the
using analysis item centering ,levelling or centering/leveling.Once workpiece is alligned with
respect to the worktable ,the analysis item such as circularity or cylindricity is activated by
clicking the respective icon and as per the instructions avaialble,the roundness and cylindricity
values are calculated.
For detailed description ,the Roundness Tester automatic can be referred.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus, (Part No
b:SA-145),Form Measuring Software supplied with the machine.
Calibration: | The annual calibration of the machine is subject to satisfactory performance of the machine.
In case of any doubt in the overall performance of the machine,the periodic calibration is done.
| There are three types of calibration : |  | 1) Standard magnification Speciman
| 2) Dynamic Calibration |  | 2) Point Gauge Master
3) Diameter measurement calibration
Using a Standard magnification Speciman for Probe sensitivity calibration.
Using a Point Gauge Master for R counter reset or Z counter reset.
Diameter calibration are usually done by the maker in case of doubted performance with the
specified masters.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Documents:
Operation Standard "Annexure to ISW-QA-041"
Periodic Calibration of Roundness Tester (RA-2000DH) - "ISW-QA-041A"
page 1 of 2
Permissible Deviations for the Roundness Tester (Kosaka) :
Sr No | Test Characteristics | Test Method |  |  | Acceptable Limits of Error (µm)
1 | Roundness | Measuring by the resting the Hemispere on rotatery table & measure the roundness of Hemispere. |  |  | Roundness of the hemispear master shoule be with in 0.05µm
2 | Straightness | Measuring by the resting the Cylinderical master on rotatry table and measure the Straightness of Cylinderical master. |  |  | 0.2µm/100mm(0.8µm/600mm)
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | Precision Products Marketing |  | Roundness Tester (Kosaka)
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Bore Scope
|  | Doc No : | ISW-QA-042
|  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  | Rev No : | 0
|  | Date : | 01.04.2015
Bore Scope
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of Cam Shaft ID Burr, Dust & Rust.
Reference Standards/Materials :
Visual Limit Samples.
Description: | This Machine could be used effectively for checking of various cam shaft oil hole burr, rust & Dust.
The detailed descriptions of the various menus can be referred from Operating Manual.
Calibration: | The annual calibration of the instrument is subject to satisfactory performance of the machine.
The master Cam piece are checked for Cam Profile on Periodic Basis and the results are compared
with the Master Graph taken during the Machine Installation time.
Calibration Frequecy:
In case of suspected performance ,the instrument is calibrated by the maker .
Referred Document :
Operation Standard "Annexure to ISW-QA-042"
Checked By
Prepared By
Approved By
Released By

## Barkhousan Testing Machine
|  | Doc No : | ISW-QA-043
|  | Issue No : | 6
Musashi Auto Parts India Pvt Ltd |  | Rev No : | 0
|  | Date : | 01.04.2015
Barkhosen Testing Machine
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of Cam Shaft lob stress after grinding.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus,
Calibration Master & Software supplied with the machine.
Description: | This Machine could be used effectively for checking of various cam shaft Lob stress generated after
grinding. The components are mounter between two cneters & roated by manual mode or auto mode
as per requirment. Stylus remains in contact of lob during inspection and send the feed back to desplay
unit. After calculation at software machine provides the numeric value of stress on display.
After inspection machine provides the stress graph and then operator mark the maximum place of
stress value.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus,
Calibration Master.
Calibration: | The annual calibration of the instrument is subject to satisfactory performance of the machine.
The master Cam piece are checked for Cam stress on Periodic Basis and the results are compared
with the Master Graph taken during the Machine Installation time.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Document :
Operation Standard "Annexure to ISW-QA-043"
Checked By
Prepared By
Approved By
Released By

## Optical Shaft Measurment.
|  |  | Doc No : |  | ISW-QA-044
|  |  | Issue No : |  | 6
Musashi Auto Parts India Pvt Ltd |  |  | Rev No : |  | 0
|  |  | Date : |  | 01.04.2015
Optical Shaft Measurment Machine
Purpose: | This documented text aims to demonstrate the various features available in the abovesaid equipment,
operating procedures,calibration methods and reference standards used ,if any.
Scope: | For checking of dimensional measurements such as angle,radius and distance measurements
of various machined components.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus,
Calibration Master & Software supplied with the machine.
Description: | This machine makes use of Formpack software for measurement and evaluation of various
dimensional parameters viz angle,radius,linear dimension etc .The components are mounted between
two centers and clamp prperly. After clamping the job door of the machine has been closed. As per
loaded component programme has been selected and run the same.
Standard Accessories :
The standard accessries required for the operation of the machine are Standard Stylus,
Calibration Master.
Calibration: | The annual calibration of the instrument is subject to satisfactory performance of the machine.
The master Cam piece are checked for Cam stress on Periodic Basis and the results are compared
with the Master Graph taken during the Machine Installation time.
Calibration Frequecy:
Once in a Year & In case of suspected performance ,the machine is calibrated by the maker .
Referred Document :
Operation Standard "Annexure to ISW-QA-044"
page 1 of 2
Permissible Deviations for the Optical Shaft Measurement :
Sr No | Test Characteristics | Test Method |  | Acceptable Limits of Error (µm)
1 | Lengths | Measuring by the putting the Step Master between thr upper & lower center and clamp properly. Close the door & open Softwear & measured various lengths of step master. |  | L=(5+L[mm]/100)µm
|  |  |  | 50 mm | ـ | 5.5µm
|  |  |  | 100mm | ـ | 6µm
|  |  |  | 200mm | ـ | 7µm
|  |  |  | 300mm | ـ | 8µm
|  |  |  | 400mm | ـ | 9µm
|  |  |  | 500mm | ـ | 10µm
2 | Diameters | Measuring by the putting the Step Master between thr upper & lower center and clamp properly. Close the door & open Softwear & measured various Diameters of step master. |  | D=(2+D[mm]/100)µm
|  |  |  | 10mm | ـ | 2.1µm
|  |  |  | 15mm | ـ | 2.15µm
|  |  |  | 20mm | ـ | 2.2µm
|  |  |  | 25mm | ـ | 2.25µm
|  |  |  | 30mm | ـ | 2.3µm
|  |  |  | 45mm | ـ | 2.35µm
|  |  |  | 50mm | ـ | 2.4µm
Identification display method:
Calibration sticker on machine provided by maker or MAP-ID.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Tracibility:
| NABL | JENOPTIK | Optical Shaft Measurement
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Master Assembly
[5 empty] | Doc No : | ISW-QA-045
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-07-21
Calibration of Master Assembly
Calibration Method: |  |  | ISW-QA-045
Calibration Frequency: |  |  | As per Calibration Frequency Chart.
1.0 Scope: |  | This method describes the making and calibration of Master Assembly.
2.0 References : |  | Child Part Drawing.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min).
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Disc Micrometer, Dial Indicator (0.001 mm) & PCD Tester.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 0f 2
Permissible Deviations for the Master Assembly :
Sl No | Measuring Item |  | Measuring Method |  | Acceptable limits of Error
1 | Span Size |  | Ensure that the instrument has been made to read zero at the minimum measuring length. Now check the span size of child part between no of teeths as specified in drawing and observe the value. Repeat the procedure for all the Child Parts and note down the observed values. |  | For New Making | 50 % of the tolerence specified in Part Drawing
2 | PCD R/O |  | Put the Child Part on taper Mandrel. Next mount the Child Part on PCD Tester. Mount the PCD Pin (size as given in drawing or process sheet) in PCD Tester. Now adjust the PCD pin on a teeth of Child Part and set the Dial Indicator at zero position. Now rotate the part by one teeth so that PCD Pin rests on next teeth. Repeat the same procedure and complete one revolution and observe the total deviation in Dial Indicator. |  | For New Making | 50 % of the tolerence specified in Part Drawing
Identification display method:
White Colour for Master Assembly Identification.
Calibration Due Date itching on Master Assembly as per Method provided by MAP-ID.
Recorded Document:
Calibration Record of Master Assembly in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCK | DISC MICROMETER
Checked By
Prepared By
Approved By
Released By
MASTER ASSEMBLY
Checked By
Prepared By
Approved By
Released By

## Surface Plate
[5 empty] | Doc No : | ISW-QA-046
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-07-28
Calibration of Surface Plate
Calibration Frequency: |  | 1 Year
1.0 Scope: |  | This method describes the calibration of Surface Plates.
2.0 References : |  | This document is based on JIS B 7513
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Enviorment :
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Calibration: |  | To ensure the performance, the instrument is calibrated by External Calibration Agency
Recorded Document:
Calibration cerfificate given by the Calibration Agency after calibration.
Calibration Record of Surface Plate in GCS.
Identification display method:
Calibration sticker on Surface Plate provided by Calibration Agency.
Calibration Sticker provided by Calibration Agency
Prepared By |  | Checked By |  | Approved By |  | Released By
page 1 0f 2
Permissible Deviations for Surface Plate :
Sr No | Test Characterstics |  | Range | Length of Diagonal Line | Acceptable limits of Error (µm)
[5 empty] | Grade 0 | Grade 1 | Grade 2
1 | Flatness |  | 160X100 | 188 | 3 | 6 | 12
|  |  | 250X160 | 296 | 3.5 | 7 | 14
|  |  | 250X250 | 354 | 3.5 | 7 | 15
|  |  | 400X250 | 471 | 4 | 8 | 16
|  |  | 400X400 | 566 | 4.5 | 9 | 17
|  |  | 630X400 | 745 | 5 | 10 | 20
|  |  | 630X630 | 891 | 5 | 10 | 21
|  |  | 1000X630 | 1180 | 6 | 12 | 24
|  |  | 1000X1000 | 1414 | 7 | 14 | 28
* Length of Diagonal Line or Informative Reference :
In the case where the size of the working surface is different in size from the nominal
size, the tolerence value for the flatness shall be calculated in accordance with length
of diagonal line (Informative Reference).
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | SURFACE PLATE
Prepared By |  | Checked By |  | Approved By |  | Released By

## Master Gear
[5 empty] | Doc No : |  | ISW-QA-047
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | Issue No : | 1
[5 empty] | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-08-04
Calibration of Master Gear
Calibration Method: |  |  | ISW-QA-047
Calibration Frequency: |  |  | 3 months & 6 months
1.0 Scope: |  | This method describes the calibration of Master Gear
2.0 References : |  | Master Gear Drawing.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Gear Tooth Micrometer, CNC Gear Tester, CMM.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 2 to 5 observations are taken for the particular parameter.
page 1 of 3
Permissible Deviations for Master Gear :
Sr No | Test Characteristics |  | Measuring Instrument |  | Test Method |  |  | Acceptable Limit
1 | Over Ball Dia |  | Gear Tooth Mic & Ball Anvil |  | Mount the Ball Anvil (size as mentioned in drawing) in Gear Tooth Micrometer. Ensure the instrument has been made to read zero at the minimum measuring length. Now prepare two blocks with slip gauges with a diffrenece of 0.825 mm. Next keep the Master Gear on surface plate and place the Gear Tooth Micrometer at the teeth of Master Gear. Rest both the anvils of Gear Tooth Mic on the slip blocks with left anvil resting on block of lesser width and turn the ratchet to observe the value of O.B.D. |  |  | As per Drawing
2 | Lead - Profile Error & Pitch Error |  | CNC Gear Tester | Allow an optical flat to contact closely with the measuring face & take the reading of number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. | Put the Master Gear on Taper Mandrel and then mount on CNC Gear Tester. Now check the face runout of the Master Gear and if found more than 0.005 mm then true the face runout for less than 0.005 mm. Next select the program of the Master Gear and position the Stylus near the teeth of Master Gear and run the program to obtain the Report for Lead, Profile and Pitch Error. |  | 1.0 Micron | As per Drawing
25< MR£50
50< MR£ 75
75< MR£100
3 | OD Runout |  | CNC Gear Tester & Dial Test Indicator(0.001 mm) | Allow an optical flat to contact closely with the measuring face & take the reading of number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. | After inspection of Lead & Profile Error mount the magnetic dial stand on CNC Gear Tester and adjust the stylus of Dial Test Indicator on the tip of Teeth of Master Gear and set the Dial Test Indicaor at zero position. Next rotate the Master Gear to complete one revolution and obtain the total deviation in Dial Test Indicator |  | 1.0 Micron | As per Drawing
25< MR£50
50< MR£ 75
75< MR£100
4 | O.D. |  | Outside Micrometer (for even no. Of teeth) |  | Make ensure the instrument has been made to read zero at minimum measuring length. Place the micrometer on opposite teeth of master gear, turn the ratchet and obtain the value of O.D. |  |  | As per Drawing
|  |  | CMM (for Odd no. Of Teeth) |  | Place the Master Gear on CMM. Next place the stylus on the O.D of Master Gear and trace the O.D. to obtain the value.
page 2 of 3
Identification display method:
Colour Paint Marker identification on Master Gear as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Master Gear in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | GEAR TOOTH MICROMETER
MASTER GEAR
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Blade Micrometer
[5 empty] | Doc No : | ISW-QA-048
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-08-04
Calibration of Blade Micrometer
Calibration Method: |  |  | ISW-QA-048
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Blade Micrometer for External Measurement.
2.0 References : |  | This document is based on JIS B 7502
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Micrometer Checker, Slip Gauge Grade-0, Comparator Stand and Dial Gauge (0.001).
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for Blade Micrometer :
Blade Micrometer for External Measurement:
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC :- | 0.001 mm
1 | Instrumental Error Verify at 2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80 & 25 mm combination of Gauge Blocks |  | Ensure the instrument has been made to read zero at the minimum measuring length. Next, insert a gauge block between the measuring faces,turn the ratchet and obtain the difference between the reading of Blade Micrometer and the dimension of the gauge block . |  | 0 < MR £25 | ± 2 microns
25 < MR £50
50 < MR £ 75
[5 empty] | 75 < MR £100 | ± 3 microns
100 < MR £125
125 < MR £150
[5 empty] | 150 < MR £175 | ± 4 microns
2 | Graduations |  | Visual |  | Good Contrast
3 | Thread Free Play |  | Visual |  | No perceptible free play (adjust threads if required)
4 | Constant Force Mechanism (if applicable) |  | Visual |  | Uniform Pressure
Identification display method:
Colour cable tie identification on Blade Micrometer as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Blade Micrometer in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | MICROMETER CHECKER
BLADE MICROMETER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Gear Tooth Micrometer
[5 empty] | Doc No : | ISW-QA-049
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-08-04
Calibration of Gear Tooth Micrometer
Calibration Method: |  |  | ISW-QA-049
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Gear Tooth Micrometer for External Measurement.
2.0 References : |  | This document is based on JIS B 7502
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Micrometer Checker, Slip Gauge Block Grade 0 & Round Anvil.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for Gear Tooth Micrometer :
Gear Tooth Micrometer for External Measurement:
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC :- | 0.001 mm
1 | Instrumental Error Verify at 2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80 & 25 mm combination of Gauge Blocks |  | Clamp the round anvils in Gear Tooth Micromter. Ensure the instrument has been made to read zero at the minimum measuring length.Next,insert a gauge block between the measuring faces,turn the ratchet and obtain the difference between the reading of Gear Tooth Micrometer and the dimension of the gauge block |  | 0 < MR £ 25 | ± 2 microns
25 < MR £ 50
50 < MR £ 75
[5 empty] | 75 < MR £ 100 | ± 3 microns
100 < MR £ 125
125 < MR £ 150
[5 empty] | 150 < MR £175 | ± 4 microns
175 < MR £ 200
200 < MR £ 225
2 | Graduations |  | Visual |  | Good Contrast
3 | Thread Free Play |  | Visual |  | No perceptible free play (adjust threads if required)
4 | Flat on Round Anvil |  | No visible flats |  | Magnifying Glass or microscopes
5 | Constant Force Mechanism (if applicable) |  | Visual |  | Uniform Pressure
Identification display method:
Colour cable tie identification on Gear Tooth Micrometer as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of External Micrometer in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | MICROMETER CHECKER/ GAUGE BLOCK
GEAR TOOTH MICROMETER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Point Micrometer
[5 empty] | Doc No : | ISW-QA-050
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-08-04
Calibration of Point Micrometer
Calibration Method: |  |  | ISW-QA-050
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Point Micrometer for External Measurement.
2.0 References : |  | This document is based on JIS B 7502
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Micrometer Checker, Slip Gauge Block Grade- 0
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for the Point Micrometer :
Point Micrometer for External Measurement:
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC :- | 0.001 mm
1 | Instrumental Error Verify at 2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80 & 25 mm combination of Gauge Blocks |  | Ensure the instrument has been made to read zero at the minimum measuring length.Next,insert a gauge block between the measuring faces,turn the ratchet and obtain the difference between the reading of micrometer callipers and the dimension of the gauge block. |  | 0 < MR £25 | ± 2 microns
25 < MR £50
50 < MR £ 75
[5 empty] | 75 < MR £100 | ± 3 microns
100 < MR £125
125 < MR £150
[5 empty] | 150 < MR £175 | ± 4 microns
2 | Graduations |  | Visual |  | Good Contrast
3 | Thread Free Play |  | Visual |  | No perceptible free play (adjust threads if required)
4 | Constant Force Mechanism (if applicable) |  | Visual |  | Uniform Pressure
Identification display method:
Colour cable tie identification on Point Micrometer as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of External Micrometer in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | MICROMETER CHECKER/ GAUGE BLOCK
POINT MICROMETER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Depth Micrometer
[5 empty] | Doc No : | ISW-QA-051
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-08-04
Calibration of Depth Micrometer
Calibration Method: |  | 0
Calibration Frequency: |  | 3 months
1.0 Scope: |  | This method describes the calibration of Depth Micrometers for Distance Measurement.
2.0 References : |  | This document is based on JIS B 7544
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Slip Gauge Block Grade 0 or Grade 1, Surface Plate of Grade 0 or Grade 1 &
Optical Flat Grade 1 or Grade 2.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
Prepared By |  | Checked By |  | Approved By |  | Released By
page 1 of 3
Permissible Deviations for Depth Micrometer :
Sr No | Test Characteristics |  | Test Method |  |  | Acceptable Limits of Error
[6 empty] | LC:- | 0.001 mm
1 | Instrumental Error Verify at 2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80 & 25 mm combination of Gauge Blocks |  | Ensure the instrument has been made to read zero at the minimum measuring length. Place a pair of gauge blocks (equal length) of longer than the maximum measuring length on a precision surface plate, and allow the reference plane of the micrometer to contact closely to this. Next place the gauge block between the measuring plane of micrometer and the surface plate, rotate the ratchet of the micrometer and obtain the difference between the reading of micrometer and the dimension of the gauge block. |  |  | 0 < MR £ 25 | ± 4.0 Micron
[6 empty] | 25 < MR £ 50 | ± 5.0 Micron
[6 empty] | 50 < MR £ 75 | ± 6.0 Micron
75 < MR £ 100
[6 empty] | 100 < MR £ 150 | ± 7.0 Micron
2 | Flatness of Measuring Plane |  | Allow the optical flat to contact closely to the measuring plane & read out the number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. |  |  | 0 < MR £ 25 | 0.60 Micron
25 < MR £ 50
50 < MR £ 75
75 < MR £ 100
100 < MR £ 150
3 | Flatness of Reference Plane |  | Allow the optical flat to contact closely to the reference palne & read out the number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. |  |  | 0 < MR £ 25 | 2.0 Micron
25 < MR £ 50
50 < MR £ 75
75 < MR £ 100
100 < MR £ 150
4 | Parallelism of Measuring Plane with respect to Reference Plane |  | Fix a pair of gauge blocks of equal length on a surface plate, and allow the reference plan of micrometer to contact closely to these. Next insert a gauge block between the surface plate and 4 corners of measuring plane of micrometer in turn, rotate the ratchet of the micrometer to read the respective readings and obtain its maximum difference. |  |  | 0 < MR £ 25 | 5.0 Micron
25 < MR £ 50
[6 empty] | 50 < MR £ 75 | 6.0 Micron
75 < MR £ 100
[6 empty] | 100 < MR £ 150 | 7.0 Micron
5 | Graduations |  | Visual |  |  | Good Contrast
6 | Thread Free Play |  | Visual |  |  | No perceptible free play (adjust threads if required)
7 | Constant Force Mechanism (if applicable) |  | Visual |  |  | Uniform Pressure
Prepared By |  | Checked By |  | Approved By |  | Released By
page 2 of 3
Identification display method:
Colour Paint Marker identification on Depth Micrometer as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Depth Micrometer in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  | Released By
NABL
EXTERNAL CALIBRATION AGENCY
GAUGE BLOCK
DEPTH MICROMETER

## Inside Groove Vernier
[5 empty] | Doc No : | ISW-QA-052
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-08-04
Calibration of Inside Groove Vernier Calliper
Calibration Method: |  | ISW-QA-052
Calibration Frequency: |  | 3 months
1.0 Scope: |  | This method describes the calibration of Inside Groove Vernier Calliper for Internal Measurement.
2.0 References : |  | This document is based on JIS B 7507
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Caliper Checker
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
Prepared By |  | Checked By |  | Approved By |  | Released By
page 1 of 2
Permissible Deviations for Inside Groove Vernier :
Inside Groove Vernier for Internal Measurement.
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC:- | 0.05 mm | 0.02 or 0.01 mm
1 | Verify Instrumental Error at 20, 50, 100, 150, 200 & 300 |  | Ensure the instrument has been made to read zero at minimum measuring distance. Next measure the internal dimensions of the blocks of Calliper Checker and obtain difference between dimension of the Caliper Checker and the reading of the Inside Groove Vernier Caliper. |  | ML ≤ 50 | ± 0.050 | ± 0.020
[5 empty] | 50 < ML £ 100 | ± 0.060 | ± 0.030
[5 empty] | 100 < ML £ 200 | ± 0.070
[5 empty] | 200 < ML £ 300 | ± 0.080 | ± 0.040
[5 empty] | 300 < ML £ 400 | ± 0.090
[5 empty] | 400< ML £ 500 | ± 0.100 | ± 0.050
[5 empty] | 500< ML £ 600 | ± 0.110
[5 empty] | 600< ML £700 | ± 0.120 | ± 0.060
[5 empty] | 700< ML £800 | ± 0.130
[5 empty] | 800< ML £900 | ± 0.140 | ± 0.070
[5 empty] | 900< ML £1000 | ± 0.150
Identification display method:
Colour Paint Marker identification on Inside Groove Vernier Calliper as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Inside Groove Vernier Calliper in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  | Released By
NABL
EXTERNAL CALIBRATION AGENCY
CALLIPER
CHECKER
INSIDE GROOVE VERNIER CALLIPER

## Depth Vernier
[5 empty] | Doc No : | ISW-QA-053
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-08-04
Calibration of Depth Vernier
Calibration Method: |  | ISW-QA-053
Calibration Frequency: |  | 3 months
1.0 Scope: |  | This method describes the calibration of Depth Verniers for Distance Measurement.
2.0 References : |  | This document is based on JIS B 7518
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Slip Gauge Block Grade 2, Precision Surface Plate of Grade 1 , Measuring Device &
Lever type Dial Gauge 0.002 mm in scale.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
Prepared By |  | Checked By |  | Approved By |  | Released By
page 1 of 3
Permissible Deviations for Depth Vernier :
Depth Vernier for Distance Measurement:
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC :- | 0.05 mm | 0.02 or 0.01 mm
1 | Verify Instrumental Error at 20, 50, 100, 150, 200 & 300 |  | Ensure the instrument has been made to read zero at minimum measuring distance. Under the condition where the measuring surface of base has been allowed to contact closely to two pieces of block gauges equal in nominal size on a precision surface plate, obtain the instrumental error by subtracting the dimension of gauge block from reading of the depth vernier, when the measuring surface of main scale has been applied to the surface plate. |  | ML ≤ 50 | ± 0.05 | ± 0.02
[5 empty] | 50 < ML £ 100 | ± 0.06 | ± 0.03
[5 empty] | 100 < ML £ 200 | ± 0.07
[5 empty] | 200 < ML £ 300 | ± 0.08 | ± 0.04
2 | Flatness for Measuring Surface of base |  | Preliminarily adjust so that the measuring surface of the base and measuring surface of main scale come on the same plane approx., attach it to the measuring device after has been fixed with the locking screw, carry out inclination adjustment so that the readings of lever type dial gauge applied to both ends of measuring surface of base become equal, then measure by travelling the lever type dial gauge in the left and right on the measuring surface of base, and obtain difference between the maximum and minimum value acquired. |  | ML >= 50 | 5 Micron
50 < ML £ 100
100 < ML £ 200
200 < ML £ 300
3 | Parallelism for Measuring Surface of main scale with respect to Measuring Surface of base |  | In succession to the measurement of no. 2, apply the measuring pointer of lever type dial gauge to the measuring surface of main scale, measure by travelling in the left and right, and obtain the difference between the maximum and minimum value acquired. |  | ML >= 50 | 5 Micron
50 < ML £ 100
100 < ML £ 200
200 < ML £ 300
4 | Graduations |  | Visual |  | Good Contrast
5 | Thread Free Play |  | Visual |  | No perceptible free play (adjust threads if required)
6 | Constant Force Mechanism (if applicable) |  | Visual |  | Uniform Pressure
Prepared By |  | Checked By |  | Approved By |  | Released By
page 2 of 3
Identification display method:
Colour Paint Marker identification on Depth Vernier as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Depth Vernier in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  | Released By
NABL
EXTERNAL CALIBRATION AGENCY
GAUGE BLOCK
DEPTH VERNIER

## Pasa Meter
[5 empty] | Doc No : | ISW-QA-054
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-08-04
Calibration of Pasa Meter
Calibration Method: |  |  | ISW-QA-054
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Pasa Meter for External Measurement.
2.0 References : |  | This document is based on JIS B 7520
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Slip Gauge Block Grade 0 ,Optical Flat Grade 1 or Grade 2
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 3
Permissible Deviations for Pasa Master :
Pasa Meter for External Measurement:
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
[5 empty] | LC :- | 0.001 mm
1 | Instrumental Error Verify at 15.00 mm combination of Gauge Blocks |  | Insert the gauge block between measuring face. Adjust the pointer at zero position under loaded condition, now introduce the gauge block with addition & subtraction of 10 microns respectively and obtain the diffrence between reading of pasameter and dimension of gauge block. Carry out this procedure throughout the entire range of indicating dial at intervals of 10 microns and take the difference in reading of pasameter and dimension of gauge block. |  | 0< MR£ 25 | ± 2 microns
25< MR£50
50< MR£ 75
75< MR£100
2 | Flatness of Measuring Face |  | Allow an optical flat to contact closely with the measuring face & take the reading of number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. |  | 0< MR£25 | 1.0 Micron
25< MR£50
50< MR£ 75
75< MR£100
3 | Parallelism of Measuring Face |  | Interpose a gauge block at center between the both measuring faces and at 4 places of peripheries in turn, take readings of respective values by the pointer under the loaded condition & obtain the maximum difference thereof. |  | 0< MR£25 | 2.0 Micron
25< MR£50
50< MR£ 75
75< MR£100
4 | Thread Free Play |  | Visual |  | No perceptible free play (adjust threads if required)
5 | Chamfer on Anvil |  | Visual |  | Chamfer should be available on the edges of Anvils of Pasameter.
6 | Constant Force Mechanism (if applicable) |  | Visual |  | Uniform Pressure
page 2 of 3
Identification display method:
Colour cable tie identification on Pasa Meter as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Pasa Meter in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | GAUGE BLOCK
PASA METER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## PCD Tester
[5 empty] | Doc No : | ISW-QA-055
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-08
Calibration of PCD Tester
Calibration Method: |  | ISW-QA-055
Calibration Frequency: |  | 1 year
Calibration: |  | To ensure the performance, the instrument is calibrated by the External Agency.
1.0 Scope: |  | This method describes the calibration of PCD Tester with Center Height 125 mm to
300 mm.
2.0 References : |  | This document is based on IS 5980-1978
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
Center Height (H): Height of Centers from the top face of the bed.
Center Distance: Distance between two centers.
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Test Mandrel (Straight), Magnetic Stand and Dial Test Indicator.
6.0 Calibration Process:
| 6.1 | The appearance ,cleaniness and condition of the centers(no dents, no damage and no
worn out) of the PCD Tester are verified.
| 6.2 | The Height Column and Push Handle of PCD Tester should be working properly.
| 6.3 | The Head and Tailstock should be sliding smoothly.
| 6.4 | PCD Tester is calibrated for the Coaxiality of Centers and parallelism of the Axis of
centers with respect to Guideways.
Prepared By |  | Checked By |  | Approved By |  | Released By
|  | Checked By |  |  | Approved By
page 1 of 3
Permissible Deviations for the PCD Tester :
Sr No | Test Characterstic | Method |  | Acceptence Limit of Error (mm)
|  |  |  | Center Height:- |  | 125 and 160 (mm) | 200 to 300 (mm)
1 | Co-axiality of Centers | A Cylendrical test mandrel is held between the centers. A suitable dial indicator is mounted on the slide. The slide is moved and the readings are taken in two planes perpendicular to each other close to the ends of the mandrel and also at intervals of 300 mm. The length of the mandrel shall be nearly equal to center distance. The difference in values is obtained and shall not exceed the permissible value. |  | Over any length of 300 mm between Centers |  | 0.01 | 0.015
|  |  |  | Center Distance upto 500 mm |  | 0.015 | 0.02
|  |  |  | Center Distance above 500 mm |  | 0.02 | 0.04
2 | Parallelism of the Axis of centers with respect to Guideways | A cylendrical test mandrel of 300 mm length with taper shank is inserted into the taper bore of the center. A suitable dial indicator is mounted on the base of PCD Tester. Readings are taken in two planes perpendicular to each other at two positions close to the ends of the mandrel. The difference in values is obtained and shall not exceed the permissible value. This is repeated for the other center bore and at different positions along the length of base. |  | Per 300 mm leaning towards the free end of the mandrel |  | 0.01 | 0.015
3 | General Feel of Instrument |  |  | Smooth Movement , No Stickness,No Center Damage, No Dents etc…
Prepared By |  | Checked By |  | Approved By |  | Released By
|  | Checked By |  |  | Approved By
page 2 of 3
Calibration Certificate:
The following measuring items shall be contained in the calibration certificate.
a) | Co-axiality of Centers.
b) | Parallelism of the Axis of centers with respect to Guideways.
Identification display method:
Calibration sticker on PCD Tester Provided by Calibration Agency
Calibration Sticker provided by Maker
Recorded Document:
Calibration cerfificate given by External Calibration Agency after calibration.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | PCD TESTER
Prepared By |  | Checked By |  | Approved By |  | Released By
|  | Checked By |  |  | Approved By

## Slip Gauge Block
[6 empty] | Doc No : |  | ISW-QA-056
[6 empty] | Issue No : |  | 1
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : |  | 0
[6 empty] | Date : |  | 2016-08-08
Calibration of Slip Gauge Blocks
Calibration Method: |  |  | ISW-QA-056
Calibration Frequency: |  |  | 12 months
Calibration: |  | To ensure the performance, the instrument is calibrated by the External Agency.
1.0 Scope: |  | This method describes the procedure for calibration of Slip Gauge Blocks.
2.0 References : |  | This document is based on JIS B 7506 : 1997
3.0 Description: |  | A Gauge Block is an end standard of rectangular section, made of wear resistant material,
with one pair of plane, mutually parallel meausering faces. The measuring faces shall have
surfaces which can be wrung to the measuring faces of other gauge blocks or plane.
4.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
Deviation of Length (e): The deviation of length at any point of gauge block from nominal length.
Variation in Length (v): The difference between the maximum length and minimum length of
a gauge block.
Wringing: Wringing is the property of the measuring faces of gauge block to adhere completely
to other measuring faces or to faces with similar surface finish as a result of molecular forces.
5.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
6.0 Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
7.0 Grades: |  | The Gauge Block shall be classified according to accuracy into four grades of grade K,
grade 0, grade 1 and grade 2.
Prepared By |  |  | Checked By |  | Approved By |  |  | Released By
page 1 of 3
Permissible Deviations for Slip Gauge Blocks :
1) Tolerence for Limit Deviation of length at any point & tolerence for variation in length:
Nominal Length Range (mm) |  | Grade K |  | Grade 0 |  | Grade 1 |  | Grade 2
|  | te (µm) | tv (µm) | te (µm) | tv (µm) | te (µm) | tv (µm) | te (µm) | tv (µm)
0.50≤MR≤10 |  | ±0.20 | 0.05 | ±0.12 | 0.1 | ±0.20 | 0.16 | ±0.45 | 0.3
10<MR≤25 |  | ±0.30 | 0.05 | ±0.14 | 0.1 | ±0.30 | 0.16 | ±0.60 | 0.3
25<MR≤50 |  | ±0.40 | 0.06 | ±0.20 | 0.1 | ±0.40 | 0.18 | ±0.80 | 0.3
50<MR≤75 |  | ±0.50 | 0.06 | ±0.25 | 0.12 | ±0.50 | 0.18 | ±1.00 | 0.35
75<MR≤100 |  | ±0.60 | 0.07 | ±0.30 | 0.12 | ±0.60 | 0.2 | ±1.20 | 0.35
100<MR≤150 |  | ±0.80 | 0.08 | ±0.40 | 0.14 | ±0.80 | 0.2 | ±1.60 | 0.4
150<MR≤200 |  | ±1.00 | 0.09 | ±0.50 | 0.16 | ±1.00 | 0.25 | ±2.00 | 0.4
120<MR≤250 |  | ±1.20 | 0.1 | ±0.60 | 0.16 | ±1.20 | 0.25 | ±2.40 | 0.45
250<MR≤300 |  | ±1.40 | 0.1 | ±0.70 | 0.18 | ±1.40 | 0.25 | ±2.80 | 0.5
300<MR≤400 |  | ±1.80 | 0.12 | ±0.90 | 0.2 | ±1.80 | 0.3 | ±3.60 | 0.5
400<MR≤500 |  | ±2.20 | 0.14 | ±1.10 | 0.25 | ±2.20 | 0.35 | ±4.40 | 0.6
500<MR≤600 |  | ±2.60 | 0.16 | ±1.30 | 0.25 | ±2.60 | 0.4 | ±5.00 | 0.7
600<MR≤700 |  | ±3.00 | 0.18 | ±1.50 | 0.3 | ±3.00 | 0.45 | ±6.00 | 0.7
700<MR≤800 |  | ±3.40 | 0.2 | ±1.70 | 0.3 | ±3.40 | 0.5 | ±6.50 | 0.8
800<MR≤900 |  | ±3.80 | 0.2 | ±1.90 | 0.35 | ±3.80 | 0.5 | ±7.50 | 0.9
900<MR≤1000 |  | ±4.20 | 0.25 | ±2.00 | 0.4 | ±4.20 | 0.6 | ±8.00 | 1
2) Tolerence for Flatness of measuring face:
Nominal Length Range (mm) |  | Grade K (µm) |  | Grade 0 (µm) |  | Grade 1 (µm) |  | Grade 2 (µm)
0.50≤MR≤150 |  | 0.05 |  | 0.1 | 0.1 | 0.15 | 0.16 | 0.25 | 0.3
150<MR≤500 |  | 0.1 |  | 0.15 |  | 0.18 |  | 0.25
500<MR≤1000 |  | 0.15 | 0.05 | 0.18 | 0.1 | 0.2 | 0.16 | 0.25 | 0.3
Prepared By |  |  | Checked By |  | Approved By |  |  | Released By
page 2 of 3
Calibration Certificate :
The following measuring items shall be contained in the calibration certificate.
a) | Central length or deviation of central length from nominal length.
b) | Quoted uncertainity.
c) | Statement of traceability.
d) | Coefficient of thermal expansion.
Identification display method:
Calibration sticker on Slip Gauge Box Provided by Calibration Agency
Calibration Sticker provided by Maker
Recorded Document:
Calibration cerfificate given by External Calibration Agency after calibration.
Tracibility:
|  | NABL |  | EXTERNAL CALIBRATION AGENCY |  | SLIP GAUGE BLOCK
Prepared By |  |  | Checked By |  | Approved By |  |  | Released By

## Calliper Checker
|  |  |  | Doc No : | ISW-QA-057
|  |  |  | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  | Rev No : | 0
|  |  |  | Date : | 2016-08-10
Caliper Checker
Purpose: | This documented text aims to demonstrate the various features available in the above said equipment,
operating procedures, acceptance limits and reference standards used ,if any.
Scope: | This method describes the procedure for calibration of Vernier Calipers and Height Gauges using
Caliper Checker, and the acceptance limits for Caliper Checker.
Reference : | This document is based on Makers Catalogue.
Description: | This Equipment could be used effectively for calibration of Vernier Calipers and Height Gauges.
Caliper Checker is designed to check accuracy of Main Scale at regular intervals through out the range,
which meets the requirement of the Standards. The instrument is sufficiently rigid and consists of stepped
slip gauges permanently fixed in inner casing. The Caliper Checker also stands perpendicular to a surface
for height gauge inspection.
Calibration of Vernier Caliper:
Make the Vernier Caliper to read zero at minimum measuring dimension.Check the size of Gauge Blocks
with external jaws and internal jaws of the vernier caliper and obtain the differenece between reading of
vernier caliper and size of gauge block. The difference in value is compared with the standard specified
in JIS and if lie within specified range, Vernier Caliper is accepted.
Calibration of Height Gauge:
Keep the Caliper Checker in perpendicular position. Make the Height Gauge read zero at zero position
of the Caliper Checker. Now check the height of the Gauge Block and obtain the differenece between
reading of height gauge and height of gauge block. The difference in value is compared with the standard
specified in JIS and if lie within specified range, Height Gauge is accepted.
Technical Specifications :
| Range | Height Measurement | External Measurement |  | Internal Measurement
| 0~600 | up to 670 | up to 670 |  | up to 600
Calibration: | To ensure the performance, the instrument is calibrated by the External Calibration Agency.
Calibration Frequecy:
The annual calibration is subject to satisfactory performance of the Instrument.
Recorded Document:
Calibration cerfificate given by External Calibration Agency after calibration.
Calibration Certificate :
The following measuring items shall be contained in the calibration certificate.
a) | Expanded uncertainity.
b) | Statement of traceability.
c) | Nominal value, observed value and error of each step of Calliper Checker.
Prepared By |  | Checked By | Approved By |  | Released By
page 1 of 2
Permissible Deviations for Caliper Checker :
Sr No | Test Characteristics |  | Equipment Used |  | Acceptable Limits of Error
1 | Step Size for External Measurement/ Height Measurement |  | Caliper Checker, Lever Dial & Surface Plate |  | ±0.001
2 | Step Size for Internal Measurement |  | Caliper Checker, Lever Dial & Surface Plate |  | ±0.001
Identification display method:
Calibration sticker on Caliper Checker Provided by Calibration Agency or
Coloured Tepra Tape identification on Caliper Checker as per method provided by MAP-ID.
Calibration Sticker provided by Maker
Colour Coding
Tracibility:
|  | NABL | EXTERNAL CALIBRATION AGENCY |  | CALIPER CHECKER
Prepared By |  | Checked By | Approved By |  | Released By

## Micrometer Checker
[5 empty] | Doc No : | ISW-QA-058
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-17
Micrometer Checker
Purpose: | This documented text aims to demonstrate the various features available in the above said equipment,
operating procedures, acceptance limits and reference standards used ,if any.
Scope: | This method describes the procedure for calibration of all types of External Micrometers using
Micrometer Checker, and the acceptance limits for Micrometer Checker.
Reference : | This document is based on Makers Catalogue.
Description: | This Equipment could be used effectively for calibration of all types of External micrometers.
All gauge blocks are made of Alloy steel, hardened and sub zero treated for dimensional stability and
machine lapped housed in rigid housing for ease of calibration. The sizes in the set are so selected
that they ensure checking of micrometers at various radial positions. Thus in addition to the pitch
error, errors due to drunken threads also can be located.
Check the size of Gauge Blocks with the micrometer to be calibrated and observe the difference b/w
the reading of micrometer and the size of the gauge block. The difference in value is compared with
the standard specified in JIS and if lie within specified range, Micrometer is accepted.
Technical Specifications :
Size of the Gauge Blocks in Micrometer Checker:
2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80, 25.00 mm
Calibration: | To ensure the performance, the instrument is calibrated by the External Calibration Agency.
Calibration Frequecy:
The annual calibration of the instrument is subject to satisfactory performance of the Equipment.
Recorded Document:
Calibration cerfificate given by External Calibration Agency after calibration.
Calibration Certificate :
The following measuring items shall be contained in the calibration certificate.
a) | Expanded uncertainity.
b) | Statement of traceability.
c) | Nominal value, observed value and error of each Gauge Block of Micrometer Checker.
Prepared By |  | Checked By |  | Approved By |  | Released By
page 1 of 2
Permissible Deviations for Micrometer Checker :
Sr No | Test Characteristics |  |  | Equipment Used |  | Acceptable Limits of Error
1 | Step Size 2.50, 5.10, 7.70, 10.30, 12.90, 15.00, 17.60, 20.20, 22.80, 25.00 mm |  |  | Caliper Checker, Lever Dial & Surface Plate |  | ±0.001
Identification display method:
Calibration sticker on Caliper Checker Provided by Calibration Agency or
Coloured Tepra Tape identification on Micrometer Checker as per method provided by MAP-ID.
Calibration Sticker provided by Maker
Colour Coding
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | MICROMETER CHECKER
Prepared By |  | Checked By |  | Approved By |  | Released By

## Straight Mandrel
[5 empty] | Doc No : | ISW-QA-059
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-17
Calibration of Straight Mandrel
Calibration Frequency: |  | 12 months
1.0 Scope: |  | This method describes the calibration of Straight Mandrel
2.0 References : |  | This document is based on JIS B 7545 : 1982
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Calibration: |  | To ensure the performance, the instrument is calibrated by the maker.
Recorded Document:
Calibration cerfificate given by maker after calibration.
Identification display method:
Calibration sticker on Straight Mandrel provided by Maker.
Calibration Sticker provided by Maker
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | STRAIGHT MANDREL
Prepared By |  | Checked By |  | Approved By |  | Released By

## Air Converter
[5 empty] | Doc No : | ISW-QA-060
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-17
Calibration of Air Electronic Converter Module
Calibration Frequency: |  | 3 months
1.0 Scope: |  | This method describes the calibration of Air Electronic Converter Module.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Calibration: |  | To ensure the performance, the instrument is calibrated by the External Agency
Recorded Document:
Calibration cerfificate given by Calibration Agency after calibration.
Calibration Record of Air Electronic Converter Module in GCS.
Identification display method:
Calibration sticker on Air Electronic Converter Module provided by Calibration Agency.
Calibration Sticker provided by Calibration Agency
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | AIR ELECTRONIC CONVERTER MODULE
Prepared By |  | Checked By |  | Approved By |  | Released By

## Air Gauge Unit
[5 empty] | Doc No : | ISW-QA-061
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-17
Calibration of Air Gauge Unit
Calibration Frequency: |  | 3 months
1.0 Scope: |  | This method describes the calibration of Air Gauge Unit.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Calibration: |  | To ensure the performance, the instrument is calibrated by the External Agency
Recorded Document:
Calibration cerfificate given by Calibration Agency after calibration.
Calibration Record of Air Gauge Unit in GCS.
Identification display method:
Calibration sticker on Air Gauge Unit provided by Calibration Agency.
Calibration Sticker provided by Calibration Agency
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | AIR GAUGE UNIT
Prepared By |  | Checked By |  | Approved By |  | Released By

## Column Bar
[5 empty] | Doc No : | ISW-QA-062
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-17
Calibration of Electronic Column Bar
Calibration Frequency: |  | 3 months
1.0 Scope: |  | This method describes the calibration of Electronic Column Bar.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Calibration: |  | To ensure the performance, the instrument is calibrated by the External Agency
Recorded Document:
Calibration cerfificate given by Calibration Agency after calibration.
Calibration Record of Electronic Column Bar in GCS.
Identification display method:
Calibration sticker on Electronic Column Bar provided by Calibration Agency.
Calibration Sticker provided by Calibration Agency
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | ELECTRONIC COLUMN BAR
Prepared By |  | Checked By |  | Approved By |  | Released By

## Comparator Stand
[5 empty] | Doc No : | ISW-QA-063
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-25
Calibration of Comparator Stand
Calibration Frequency: |  | 1 Year
1.0 Scope: |  | This method describes the calibration of Comparator Stand.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Calibration: |  | To ensure the performance, the instrument is calibrated by External Calibration Agency
Recorded Document:
Calibration cerfificate given by the Calibration Agency after calibration.
Calibration Record of Comparator Stand in GCS.
Identification display method:
Calibration sticker on Comparator Stand provided by Calibration Agency.
Calibration Sticker provided by Calibration Agency
Prepared By |  | Checked By |  | Approved By |  | Released By
page 1 0f 2
Permissible Deviations for Comparator Stand :
Sl No | Test Characterstics |  | Range | Length of Diagonal Line | Acceptable limits of Error (µm)
[5 empty] | Grade 0 | Grade 1 | Grade 2
1 | Flatness |  | 160X100 | 188 | 3 | 6 | 12
|  |  | 250X160 | 296 | 3.5 | 7 | 14
|  |  | 400X250 | 471 | 4 | 8 | 16
|  |  | 250X250 | 354 | 3.5 | 7 | 15
* Length of Diagonal Line or Informative Reference :
In the case where the size of the working surface is different in size from the nominal size, the tolerence value for the
flatness shall be calculated in accordance with lengthof diagonal line (Informative Reference).
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | COMPARATOR STAND
Prepared By |  | Checked By |  | Approved By |  | Released By

## Load Cell
[5 empty] | Doc No : | ISW-QA-064
[5 empty] | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | 0
[5 empty] | Date : | 2016-08-25
Calibration of Load Cell
Calibration Frequency: |  | 3 Months
1.0 Scope: |  | This method describes the calibration of Load Cell.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Calibration: |  | To ensure the performance, the instrument is calibrated by the maker.
Permissible Deviations for Load Cell:
Deviation upto ± 1.00 % of total capacity of Load Cell.
Recorded Document: |  | Calibration cerfificate given by maker after calibration.
Identification display method:
Calibration sticker on Comparator Stand provided by Maker.
Calibration Sticker provided by Maker
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | LOAD CELL
Prepared By |  | Checked By |  | Approved By |  | Released By

## Vertical Gear Roll Tester
[6 empty] | Doc No : |  | ISW-QA-065
[6 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : | Rev No : | 0
[6 empty] | Date : | Date : | 2016-09-06
Calibration of Vertical Gear Roll Tester
Calibration Method: |  |  | ISW-QA-065
Calibration Frequency: |  |  | 1 month & 3 months
1.0 Scope: |  | This method describes the calibration of Vertical Gear Roll Tester.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Straight Mandrel, Dial Test Indicator.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 0f 2
Permissible Deviations for Vertical Gear Roll Tester :
Sl No | Measuring Item |  | Measuring Method |  |  | Measuring Instrument |  | Acceptable limits of Error
1 | Total Runout between centers |  | Mount Straight Mandrel between centers of Vertical Gear Roll Tester. Now set the stylus of dial test indicators at one end of Straight Mandrel, adjust the pointer of Dial Test Indicators to zero position then rotate Straight Mandrel to complete one revolution and obtain the total deviation. Repeat the same procedure at the center of the mandrel and other end of the mandrel. The deviation at all three positions should be less than permissible error. |  |  | Straight Mandrel, Dial Test Indicator (0.001 mm) |  | 0.005 mm
2 | Centers |  | Centers should not be wear and damaged. |  |  | Visual |  | No dents and Steps
3 | Slide Motion |  | Slide Motion should be smooth throughout its entire range. |  |  | Visual |  | Obstruction free movement
Identification display method:
Colour Tepra Tape identification of Vertical Gear Roll Tester (1 month) as per method provided by MAP-ID.
Colour Paint Mark identification of Vertical Gear Roll Tester (3 months) as per method provided by MAP-ID.
Colour Coding (1 month) |  |  |  | Colour Coding (3 month)
Recorded Document:
Calibration Record of Vertical Gear Roll Tester in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | STRAIGHT MANDREL |  | VERTICAL GRT
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Horizontal Gear Roll Tester
[5 empty] | Doc No : |  | ISW-QA-066
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-06
Calibration of Horizontal Gear Roll Tester
Calibration Method: |  |  | ISW-QA-066
Calibration Frequency: |  |  | 1 month & 3 months
1.0 Scope: |  | This method describes the calibration of Horizontal Gear Roll Tester.
2.0 References : |  | Gauge Drawing
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Outside Micrometer.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | Carry out the inspection of the gauge as per the drawing provided for the gauge.
| 6.6 | The parameters to be checked will be highlighted in the drawing.
| 6.7 | The acceptability of the Horizontal Gear Roll Tester is as per the specified tolerance in
the drawing.
page 1 0f 2
Permissible Deviations for Horizontal Gear Roll Tester :
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument |  | Acceptable limits of Error
1 | Locator Diameter |  | Ensure the instrument has been made to read zero at the minimum measuring length. Next,check the diameter of the Locator (as highlighted in the drawing) and obtain the value of the diameter. |  | Outside Micrometer |  | As specified in Drawing
Identification display method:
Colour Tepra Tape identification of Horizontal Gear Roll Tester (1 month) as per method provided by MAP-ID.
Colour Paint Marker identification of Horizontal Gear Roll Tester (3 months) as per method provided by MAP-ID.
Tepra Tape for 1 month frequency
Paint Marker for 3 months frequency
Recorded Document:
Calibration Record of Horizontal Gear Roll Tester in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | GAUGE BLOCK |  | OUTSIDE MICROMETER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Bend & Twist FIxture
[5 empty] | Doc No : |  | ISW-QA-067
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-06
Calibration of Bend & Twist Fixture
Calibration Method: |  |  | ISW-QA-067
Calibration Frequency: |  |  | 3 month
1.0 Scope: |  | This method describes the calibration of Bend & Twist Fixture.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
CMM, Setting Masters of Bend & Twist fixture.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 0f 2
Permissible Deviations for Bend & Twist Fixture :
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument |  | Acceptable limits of Error
1 | Bend for Setting Master |  | Place the Setting Master on Magnetic V-Block along X-axis on CMM machine. Create a plane on resting face of V-Block, and trace the Big End Diameter and Small End Diameter and create a cylinder. Now check the parallelism of two cylinders in X-Z plane at a height specified on master and observe the value for bend. |  | CMM |  | ±0.001
2 | Twist for Setting Master |  | After Bend inspection now check the parallelism of two cylinders in Y-Z plane at a height specified on master and observe the value for twist |  | CMM |  | ±0.001
Identification display method:
Colour Paint Marker identification of Bend & Twist Fixture as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Bend & Twist Fixture in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | CMM |  | BEND & TWIST FIXTURE
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Functional Gauges
[5 empty] | Doc No : |  | ISW-QA-068
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-06
Calibration of Functional Gauges
Calibration Method: |  |  | ISW-QA-068
Calibration Frequency: |  |  | 15 days, 1 month, 3 months & 6 months.
1.0 Scope: |  | This method describes the calibration of Functional Gauges.
2.0 References : |  | Gauge Drawing
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Best suited as per drawing.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | Carry out thre inspection of the gauge as per the drawing provided for the following gauge.
| 6.5 | The parameters to be checked will be highlighted in the drawing.
| 6.6 | The acceptability of the Functoinal Gauges is as per the specified tolerances in the
drawing.
page 1 0f 2
Identification display method:
Best suited among Colour Cable Tie, Tepra Tape or Paint Marker identification of the Functional Gauge will be
used as per method provided by MAP-ID.
Recorded Document:
Calibration Record of Functional Gauge in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCKS & CALLIPER CHECKER | MEASURING INSTRUMENTS |  | FUNCTIONAL GAUGES
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Mic Setting Master
[5 empty] | Doc No : |  | ISW-QA-069
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-14
Calibration of Micrometer Setting Master
Calibration Method: |  |  | ISW-QA-069
Calibration Frequency: |  |  | 6 months
1.0 Scope: |  | This method describes the calibration of Micrometer Setting Master.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Comparator Stand, Plunger Dial (0.001 mm).
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 0f 2
Permissible Deviations for Micrometer Setting Master :
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument |  | Acceptable limits of Error
1 | Master Dimension |  | Take Gauge Block of size equivalent to that of the Master. Place the Gauge Block on Comparator Stand and position the plunger dial on the top face of Gauge Block and set the plunger dial at zero position. Now remove the Gauge Block and insert the Setting Master under the plunger dial and observe the deviation in the plunger dial from zero position. |  | Comparator Stand & Plunger Dial (0.001 mm) |  | ±0.001 mm
2 | Appearence |  | Visual Inspection |  | - |  | No dents and Steps
Identification display method:
Colour Paint Marker identification of Mic Setting Master as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Micrometer Setting Master in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY | COMPARATOR STAND |  | MICROMETER SETTING MASTER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Span Master
[6 empty] | Doc No : |  | ISW-QA-070
[6 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : | Rev No : | 0
[6 empty] | Date : | Date : | 2016-09-14
Calibration of Span Master
Calibration Method: |  |  | ISW-QA-070
Calibration Frequency: |  |  | 6 months
1.0 Scope: |  | This method describes the calibration of Span Master.
2.0 References : |  | Gauge Drawing
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Caliper Micrometer.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | The acceptability of the Span Master is as per the specified tolerance in the drawing.
Prepared By |  | Checked By |  |  | Approved By |  |  | Released By
page 1 0f 2
Permissible Deviations for the Span Master :
Sr No | Measuring Item |  | Measuring Method |  |  | Measuring Instrument |  | Acceptable limits of Error
1 | Dimension X |  | Ensure the instrument has been made to read zero at the minimum measuring length. Next, insert the measuring face of span master between the measuring faces of the instrument, turn the ratchet and obtain the difference between the reading of instrument and dimension of the span master. |  |  | Caliper Micrometer |  | ±0.001 mm
2 | Parallelism of Measuring Face |  | Interpose the measuring surface of the span master between measuring face of the instrument, take readings throughout the whole length of span master by turing the rachet of the instrument & obtain the maximum difference thereof. |  |  | Outside Micrometer |  | 0.002 mm
Identification display method:
Colour Cable Tie identification on Span Master as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Span Master in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCK |  | OUTSIDE MICROMETER |  | SPAN MASTER
Prepared By |  | Checked By |  |  | Approved By |  |  | Released By

## Pasameter Master
[6 empty] | Doc No : |  | ISW-QA-071
[6 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd | [5 empty] | Rev No : | Rev No : | 0
[6 empty] | Date : | Date : | 2016-09-14
Calibration of Pasameter Master
Calibration Method: |  |  | ISW-QA-071
Calibration Frequency: |  |  | 6 months
1.0 Scope: |  | This method describes the calibration of Pasameter Master.
2.0 References : |  | Gauge Drawing
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Pasameter, Gauge Block Grade-0.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | The acceptability of the Pasameter Master is as per the specified tolerance in the
drawing.
page 1 0f 2
Permissible Deviations for the Pasameter Master :
Sr No | Measuring Item |  | Measuring Method |  |  | Measuring Instrument |  | Acceptable limits of Error
1 | Outside Diameter |  | Insert the gauge block (equivelant to the nominal size of the pasameter master) between the measuring surface of the pasameter. Adjust the pointer at zero position under loaded condition. Now introduce the pasameter master in between the measuring surface of the pasameter and obtain the difference between the reading of the pasameter and dimension of the gauge block. |  |  | Pasameter & Gauge Block |  | ±0.001 mm
2 | Cylindricity |  | Measure the 1/2 of the difference between the maximum value & the minimum value of all the diameters obtained by measurment of outside diameter as the cylindercity of gauge. |  |  | Pasameter |  | 0.002 mm
Identification display method:
Colour Cable Tie identification on Pasameter Master as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Pasameter Master in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCK |  | PASAMETER |  | PASAMETER MASTER
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Spline Arbor
[5 empty] | Doc No : |  | ISW-QA-072
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-14
Calibration of Spline Arbor
Calibration Method: |  |  | ISW-QA-072
Calibration Frequency: |  |  | 1 month & 3 month
1.0 Scope: |  | This method describes the calibration of Spline Arbor.
2.0 References : |  | Gauge Drawing
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Outside Micrometer, Cylendrical Pins and V-Mic.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | Carry out the inspection of the gauge as per the drawing provided for the following gauge.
| 6.5 | The parameters to be checked will be highlighted in the drawing.
| 6.6 | The acceptability of the Spline Arbor is as per the specified tolerance in the drawing.
page 1 0f 2
Permissible Deviations for Spline Arbor :
Sr. No. | Measuring Item |  | Measuring Method |  | Measuring Instrument |  | Acceptable limits of Error
1 | Major Diameter |  | Check the no of splines. For even no of splines use Outside Micrometer and for odd no of splines use V-Micrometer. Ensure that the instrument has been made to read zero at the minimum measuring length. Next locate the Instrument on the major dia of the spline arbor and obtain the value. |  | Outside Micrometer or V-Micrometer |  | As specified in Drawing
2 | Over Pin Diameter |  | Ensure that the instrument has been made to read zero at the minimum measuring length. Now take a set of cylendrical pin and place each cylendrical pin on opposite teeth location, and check the diameter by placing measuring surface of micrometer over cylendrical pin and observe the value. |  | Outside Micrometer, Cylendrical Pins |  | As specified in Drawing
Identification display method:
Colour Tepra Tape identification of Spline Arbor (1 month) as per method provided by MAP-ID.
Colour Cable Tie identification of Spline Arbor (3 month) as per method provided by MAP-ID.
Tepra Tape for 1 month frequency
Cable Tie for 3 month frequency
Recorded Document:
Calibration Record of Spline Arbor in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCKS | OUTSIDE MICROMETER / V-MICROMETER |  | SPLINE ARBOR
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Master Disc
[5 empty] | Doc No : |  | ISW-QA-073
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-22
Calibration of Master Disc
Calibration Method: |  |  | ISW-QA-073
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Master Disc.
2.0 References : |  | Gauge Drawing
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Height Gauge, Dial Test Indicator (0.001 mm resolution), Gauge Block Grade-0 & Bore
Dial Gauge.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | The acceptability of the Master Gear is as per the specified tolerance in the drawing.
page 1 0f 3
Permissible Deviations for the Master Disc:
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument |  | Acceptable limits of Error
1 | Outside Diameter |  | Place the Gauge Blocks (equivalent to the nominal size of the Master Disc) on surface plate. Next clamp the dial test indicator on height gauge, set the dial test indicator pointer on top face of the gauge block and set the pointer at zero on maximum dia position. Now place the Master Disk on surface plate and insert under the dial test indicator and obtain the difference between the size of Master Disc and the Gauge Block. |  | Gauge Block Grade-0, Height Gauge, Dial Test Indicator (0.001 mm) & Surface Plate Grade-0 |  | As specified in Drawing
2 | Bore Diameter |  | Set the Bore Gauge at the nominal size of the Master Disc, next insert the Bore Gauge inside the bore of Master Disc and measure the diameter in two directions perpendicular to each other in the cross section at tip and root. |  | Bore Dial Gauge |  | As specified in Drawing
3 | O.D Runout |  | Put the Master Disc on the Taper Mandrel, next mount the Master Disc with Taper Mandrel on Bench Center. Now set the indicating dial at zero position on the outside diameter of the Master Disc and slowly rotate the Master Disc to complete one rotation and obtain the total deviation on the indicating dial. |  | Taper Mandrel, Bench Center & Indicating Dial |  | As specified in Drawing
4 | Total Face Runout |  | After inspecting O.D. Runout now set the stylus of Dial Test Indicator on the face of Master Disk (at one end) and rotate the master disc to complete one rotation. Now check the runout at various points of the face from one end to other end in horizontal direction and obtain the total deviation of the indicating dial. |  | Taper Mandrel, Bench Center & Indicating Dial |  | As specified in Drawing
page 2 0f 3
Identification display method:
Colour Paint Marker identification on Master Disc as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Master Disc in GCS.
Tracibility:
NABL |  | EXTERNAL CALIBRATION AGENCY |  | GAUGE BLOCK, CALIPER CHECKER | HEIGHT GAUGE |  | MASTER DISC
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Setting Ring
[7 empty] | Doc No : |  | ISW-QA-074
[7 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd | [6 empty] | Rev No : | Rev No : | 0
[7 empty] | Date : | Date : | 2016-09-22
Calibration of Setting Ring
Calibration Method: |  |  | ISW-QA-074
Calibration Frequency: |  |  | 3 months & 6 months
1.0 Scope: |  | This method describes the calibration of Setting Ring.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Bore Dial Gauge, Gauge Block Grade A
6.0 Calibration Process:
| 6.1 | The Inspection measuring & test Equipment and Master(if any) is ensured at Standard
Temperature (20°±1°) for an hour before calibration.
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be thoroughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
| 6.6 | Dial Bore Gauge is set to read zero on gauge block combinations made.
Prepared By |  |  | Checked By |  |  | Approved By |  |  | Released By
page 1 0f 2
Permissible Deviations for the Setting Ring:
Sr No | Measuring Item |  |  | Measuring Method |  |  | Measuring Instrument |  | Acceptable limits of Error
1 | Inside Diameter |  |  | Measure the diameter in two directions perpendicular to each other in the cross section at tip and root by using Dial Bore Gauge. Exclude the portions 1 mm from both ends for the measurement. |  |  | Bore Dial Gauge |  | ±0.001 mm
2 | Circularity |  |  | Measure the difference in diameter at each cross section obtained from the measurement of inside diameter 1/2, and take the maximum of the values so obtained as the circularity of gauge. |  |  | Bore Dial Gauge |  | 0.001 mm
| Cylindercity |  |  | Measure the 1/2 of the difference between the maximum value & the minimum value of all the diameters obtained by measurment of inside diameter as the cylindercity of gauge. |  |  | Bore Dial Gauge
Identification display method:
Colour Paint Marker identification on Setting Ring as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Setting Ring in GCS.
Tracibility:
| NABL |  | EXTERNAL CALIBRATION AGENCY |  | DIAL CALIBRATION TESTER |  | BORE DIAL GAUGE & DAIL INDICATOR |  | SETTING RING
Prepared By |  |  | Checked By |  |  | Approved By |  |  | Released By

## Thread Ring Gauge
[5 empty] | Doc No : |  | ISW-QA-075
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-22
Calibration of Thread Ring Gauge
Calibration Method: |  |  | ISW-QA-075
Calibration Frequency: |  |  | 1 month
1.0 Scope: |  | This method describes the calibration of Thread Ring Gauges.
2.0 References : |  | This document is based on JIS B 0251:1998
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Wear Check Plug for GO Ring & Wear Check Plug for NOT GO Ring.
6.0 Calibration Process:
| 6.1 | The Inspection measuring & test Equipment and Master(if any) is ensured at Standard
Temperature (20°±1°) for an hour before calibration.
| 6.2 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.3 | The test instrument should be throughly cleaned.
Prepared By |  | Checked By |  | Approved By |  |  | Released By
page 1 0f 2
Permissible Deviations for the Thread Ring Gauge :
The accepatability of the gauge is subject to qualification criteria with WCP (Go & Not Go).
Sr No | Measuring Item |  | Measuring Method |  | Measuring Instrument |  | Acceptable limits of Error
1 | GO Ring Gauge |  | Screw the Wear Check Plug by hand without using the excessive force, into both ends of the GO Ring Gauge and check if WCP enters the GO Ring Gauge. |  | WCP for GO Ring |  | WCP may enter into both ends of the GO Ring Gauge, but by not more than one turn of thread. If WCP can be screwed in by more than one turn of thread, the GO Ring Gauge no longer meets the specification.
2 | NOT GO Ring Gauge |  | Screw the Wear Check Ring by hand without using the excessive force, into both ends of the NOT GO Plug Gauge and check if WCR enters the NOT GO Plug Gauge. |  | WCP for NOT GO Ring |  | WCP may enter into both ends of the NOT GO Ring Gauge, but by not more than one turn of thread. If WCP can be screwed in by more than one turn of thread, the NOT GO Ring Gauge no longer meets the specification.
Identification display method:
Colour Tepra Tape identification on Thread Ring Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Thread Ring Gauge in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  |  | Released By
NABL
EXTERNAL CALIBRATION AGENCY
WCP (GO & NOT GO)
THREAD RING GAUGE

## Thread Plug Gauge
[5 empty] | Doc No : |  | ISW-QA-076
[5 empty] | Issue No : | Issue No : | 1
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-22
Calibration of Thread Plug Gauge
Calibration Method: |  |  | ISW-QA-076
Calibration Frequency: |  |  | 1 month
1.0 Scope: |  | This method describes the calibration of Thread Plug Gauges.
2.0 References : |  | This document is based on JIS B 0251:1998
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Wear Check Ring for GO Plug & Wear Check Ring for NOT GO Plug.
6.0 Calibration Process:
| 6.1 | The Inspection measuring & test Equipment and Master(if any) is ensured at Standard
Temperature (20°±1°) for an hour before calibration.
| 6.2 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.3 | The test instrument should be throughly cleaned.
Prepared By |  | Checked By |  | Approved By |  |  | Released By
page 1 0f 2
Permissible Deviations for the Thread Plug Gauge :
The accepatability of the gauge is subject to qualification criteria with WCR (Go & Not Go).
Sr. No. | Measuring Item |  | Measuring Method |  | Measuring Instrument |  | Acceptable limits of Error
1 | GO Plug Gauge |  | Screw the Wear Check Ring by hand without using the excessive force, into both ends of the GO Plug Gauge and check if WCR enters the GO Plug Gauge. |  | WCR for GO Plug |  | WCR may enter into both ends of the GO Plug Gauge, but by not more than one turn of thread. If WCR can be screwed in by more than one turn of thread, the GO Plug Gauge no longer meets the specification.
2 | NOT GO Plug Gauge |  | Screw the Wear Check Ring by hand without using the excessive force, into both ends of the NOT GO Plug Gauge and check if WCR enters the NOT GO Plug Gauge. |  | WCR for NOT GO Plug |  | WCR may enter into both ends of the NOT GO Plug Gauge, but by not more than one turn of thread. If WCR can be screwed in by more than one turn of thread, the NOT GO Plug Gauge no longer meets the specification.
Identification display method:
Colour Tepra Tape identification on Thread Plug Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Thread Plug Gauge in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  |  | Released By
NABL
EXTERNAL CALIBRATION AGENCY
WCR (GO & NOT GO)
THREAD PLUG GAUGE

## Air Snap Gauge
[5 empty] | Doc No : |  | ISW-QA-077
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | Issue No : | 1
[5 empty] | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-22
Calibration of Air Snap Gauge
Calibration Method: |  |  | ISW-QA-077
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Air Snap Gauge.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
Balancing of Jets : The position of the Air Snap gauge in the workpiece which reduces the air
flow in one jet must create a proportionate increase in air flow from an opposing jet.This means
that airflow and characteristics of the two jets must be "BALANCED"
Centrallity : Air Snap gauge Jet diameters and nozzle drops must be on a common centerline
with the inner diameter of the Snap Gauge itself.
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Setting Masters, Read Out Unit, Pasameter & Slip Gauge Block Grade 0
6.0 Calibration Process:
| 6.1 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.2 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.3 | The test instrument should be throughly cleaned.
| 6.4 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
Prepared By |  |  | Checked By | Approved By |  |  | Released By
page 1 of 3
Permissible Deviations for the Air Snap Gauge:
Sr No | Measuring Item |  | Measuring Method |  | Meas. Instrument. |  | Acceptable Limits of Error
1 | Instrumental Error |  | Insert the Gauge Block (of same dimension as of the setting master) between the measuring surface of the pasameter. Adjust the indicating dial pointer to zero position under loaded condition. Next insert the setting master in between measuring surface of pasameter. Observe the difference between the observed value of setting master and dimension of Gauge Block. |  | Passameter, Slip Gauge Block Grade 0 |  | ±0.002 mm
2 | Nozzle Balancing Checking |  | 1) Calibrate the air gauging system using the max and min masters in a normal manner. 2) Orient the gauge with a horizontal centerline so the jets are situated vertical and located midway within the width of maximum setting master. Note the exact reading. 3) Reposition the air gauge so that the nozzles are reversed 180 degrees. Observe the difference between the indicator reading from step 2 above. |  | Air Snap Gauge, Setting Masters, Read out Unit (Digital,Dial Type) |  | Differnece up to max one division on the scale (max 1 micron) for new bought out parts. This difference could be allowed till max 3 divisions in periodic checking.
3 | Centrality Checking |  | 1) Calibrate the air gauging system using the max and min masters in a normal manner. 2) Orient the gauge with a horizontal centerline so the jets are situated horizontal and located midway within the width of maximum setting master. Note the exact reading. 3) Rotate the air gauge so that the jets are reversed 180 degrees. Observe the difference between the indicator reading from step 2 above. |  | Air Snap Gauge, Setting Masters, Read out Unit (Digital,Dial Type) |  | Differnece up to max two division on the scale (max 1 microns ) for new bought out parts. This difference could be allowed till max 4 divisions during periodic calibration & inspection.
Prepared By |  |  | Checked By | Approved By |  |  | Released By
page 2 of 3
Identification display method:
Colour Cable Tie identification on Air Snap Gauge as per method provided by MAP-ID.
Colour Cable Tie identification on Setting Master as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Air Snap Gauge in GCS.
Tracibility:
Prepared By |  |  | Checked By | Approved By |  |  | Released By
EXTERNAL CALIBRATION AGENCY
NABL
GAUGE BLOCK
PASAMETER
AIR SNAP GAUGE

## Air Ring Gauge
[5 empty] | Doc No : |  | ISW-QA-078
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | Issue No : | 1
[5 empty] | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2016-09-22
Calibration of Air Ring Gauge
Calibration Method: |  |  | ISW-QA-078
Calibration Frequency: |  |  | 3 months
1.0 Scope: |  | This method describes the calibration of Air Ring Gauge.
2.0 References : |  | This document is based on maker's catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
Balancing of Jets : The position of the Air ring gauge in the workpiece which reduces the air
flow in one jet must create a proportionate increase in air flow from an opposing jet.This means
that airflow and characteristics of the two jets must be "BALANCED"
Centrallity : Air Ring Gauge Jet diameters and nozzle drops must be on a common centerline
with the inner diameter of the Ring Gauge itself.
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Setting Masters, Read Out Unit, Pasameter & Slip Gauge Block Grade 0
6.0 Calibration Process:
| 6.1 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.2 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.3 | The test instrument should be throughly cleaned.
| 6.4 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
Prepared By |  |  | Checked By | Approved By |  |  | Released By
page 1 of 3
Permissible Deviations for the Air Ring Gauge:
Sr. No. | Measuring Item |  | Measuring Method |  | Meas. Instrument. |  | Acceptable Limits of Error
1 | Setting Master Diameter |  | Insert the Gauge Block (of same dimension as of the setting plugs) between the measuring surface of the pasameter. Adjust the indicating dial pointer to zero position under loaded condition. Next insert the setting plug in between measuring surface of pasameter. Observe the difference between the observed value of setting plug and dimension of Gauge Block |  | Passameter, Slip Gauge Block Grade 0 |  | ±0.002 mm
2 | Nozzle Balancing Checking |  | 1) Calibrate the air gauging system using the max and min masters in a normal manner. 2) Orient the gauge with a horizontal centerline so the jets are situated vertical and located midway within the width of maximum setting master. Note the exact reading. 3) Reposition the setting master so that the nozzles are reversed 180 degrees. Observe the difference between the indicator reading from step 2 above. |  | Air Ring Gauge, Setting Masters, Read out Unit (Digital,Dial Type) |  | Differnece up to max one division on the scale (max 1 micron) for new bought out parts. This difference could be allowed till max 3 divisions in periodic checking.
3 | Centrality Checking |  | 1) Calibrate the air gauging system using the max and min masters in a normal manner. 2) Orient the gauge with a horizontal centerline so the jets are situated horizontal and located midway within the width of maximum setting master. Note the exact reading. 3) Rotate the setting master so that the jets are reversed 180 degrees. Observe the difference between the indicator reading from step 2 above. |  | Air Ring Gauge, Setting Masters, Read out Unit (Digital,Dial Type) |  | Differnece up to max two division on the scale (max 2 microns) for new bought out parts. This difference could be allowed till max 3 divisions during periodic calibration & inspection.
Prepared By |  | Checked By |  | Approved By |  |  | Released By
page 2 of 3
Identification display method:
Colour Cable Tie identification on Air Ring Gauge as per method provided by MAP-ID.
Colour Cable Tie identification on Setting Master as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Air Ring Gauge in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  |  | Released By
EXTERNAL CALIBRATION AGENCY
NABL
GAUGE BLOCK
PASAMETER
AIR RING GAUGE

## Dial Snap Gauge
[5 empty] | Doc No : | ISW-QA-079
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 2016-09-22
Calibration of Dial Snap Gauge
Calibration Method: |  | ISW-QA-079
Calibration Frequency: |  | 3 months
1.0 Scope: |  | This method describes the calibration of Dial Snap Gauge for External Measurement.
2.0 References : |  | This document is based on JIS B 7520
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature : 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
| 3) | Slowly move the slider in the entire range ,and ensure that it moves smoothly through
its entire range (clean and lubricate if required)
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Slip Gauge Block Grade 0 ,Optical Flat Grade 1 or Grade 2
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
Prepared By |  | Checked By |  | Approved By |  | Released By
page 1 of 2
Permissible Deviations for Dial Snap Gauge :
Sr No | Test Characteristics |  | Test Method |  |  | Acceptable Limits of Error
[6 empty] | LC:- | 0.001 mm
1 | Instrumental Error Verify at 2.50, 12.90, & 25.00 mm combination of Gauge Blocks |  | Insert the gauge block between measuring face. Adjust the pointer at zero position under loaded condition, now introduce the gauge block with addition & subtraction of 10 microns respectively and obtain the diffrence between reading of dial snap gauge and dimension of gauge block. Carry out this procedure throughout the entire range of indicating dial at intervals of 10 microns and take the difference in reading of dial snap gauge and dimension of gauge block. |  |  | 0< MR ≤ 25 | ± 2.0 microns
25 < MR £ 50
50 < MR £ 75
75 < MR £ 100
2 | Flatness of Measuring Face |  | Allow an optical flat to contact closely with the measuring face & take the reading of number of red interference fringes produced by the white light. Remark- 1 Line of Red interference fringe shall be assumed to be 0.30 Micron. |  |  | 0< MR ≤ 25 | 1.0 Micron
25 < MR £ 50
50 < MR £ 75
75 < MR £ 100
3 | Parallelism of Measuring Face |  | Interpose a gauge block at center between the both measuring faces and at 4 places of peripheries in turn, take readings of respective values by the pointer under the loaded condition & obtain the maximum difference thereof. |  |  | 0< MR ≤ 25 | 2.0 Micron
25 < MR £ 50
50 < MR £ 75
75 < MR £ 100
4 | Thread Free Play |  | Visual |  |  | No perceptible free play (adjust thread if required)
Identification display method:
Colour cable tie identification on Dial Snap Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Dail Snap Gauge in GCS.
Tracibility:
Prepared By |  | Checked By |  | Approved By |  | Released By
NABL
EXTERNAL CALIBRATION AGENCY
GAUGE BLOCK
DIAL SNAP
GAUGE

## Magnetic Dial Stand
|  |  |  | Doc No : | ISW-QA-080
Musashi Auto Parts India Pvt Ltd |  |  |  | Issue No : | 1
|  |  |  | Rev No : | 0
|  |  |  | Date : | 2016-09-22
Calibration of Magnetic Dial Stand
Calibration Method: |  | ISW-QA-080
Calibration Frequency: |  | 6 months
1.0 Scope: |  | This method describes the calibration of Magnetic Dial Stand.
2.0 References : |  | This document is based on Makers Catalogue.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
6.0 Calibration Process:
| 6.1 | Ensure the ball joints of the magnetic stand are working properly and smoothly.
| 6.2 | Check if the ON-OFF switch offers instant mounting and dismounting of the magnetic base.
| 6.3 | Check if the hydraulic position lock system of magnetic dial stand is working properly.
Identification display method:
Colour Cable Tie identification on Magnetic Dial Stand as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Magnetic Dial Stand in GCS.
Prepared By |  | Checked By | Approved By |  | Released By

## Taper Cone Gauges
[5 empty] | Doc No : | ISW-QA-081
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | 1
[5 empty] | Rev No : | 0
[5 empty] | Date : | 13.12.2016
Calibration of Taper Cone Gauge
Calibration Method: |  |  | ISW-QA-081
Calibration Frequency: |  |  | 1 month
1.0 Scope: |  | This method describes the calibration of Taper Cone Gauge.
2.0 References : |  | This document is based on Gauge Drawing
3.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Height Gauge, Slip Gauge Blocks (Grade 1 or higher), Height Master and CMM.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
| 6.5 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for Taper Cone Gauge :
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
1 | Gauge Plain Diameter of Taper Ring |  | CMM |  | As per Gauge Drawing
2 | Angle of Taper Ring |  | CMM | CMM | As per Gauge Drawing | ± 3 microns
100 < MR £ 125
125 < MR £ 150
3 | Ring Width |  | Dig. Height Gauge & Slip Gauge Box |  | As per Gauge Drawing | ± 4 microns
175 < MR £ 200
200 < MR £ 225
4 | Gauge Plain Diameter of Taper Plug |  | CMM |  | As per Gauge Drawing
2
5 | Angle of Taper Plug |  | CMM | CMM | As per Gauge Drawing | ± 3 microns
100 < MR £ 125
125 < MR £ 150
6 | Plug Height |  | Dig. Height Gauge & Slip Gauge Box |  | As per Gauge Drawing | ± 4 microns
175 < MR £ 200
200 < MR £ 225
7 | Plug Width |  | Dig. Height Gauge & Slip Gauge Box |  | As per Gauge Drawing
225 < MR £ 250
3 | Parallelism of Measuring Face |  | Interpose a gauge block at four corners of the measuring face in turn, take readings of respective values by turning the rachet of the micrometer calipers & obtain the maximum difference thereof. |  | 0 < MR £ 25 | 2 microns
8 | Play in Assy. |  | Dig. Height Gauge |  | 0.02 mm
50 < MR £ 75
[5 empty] | 75 < MR £ 100 | 3 microns
9 | Master Height Verification |  | Standard master & Dig. Height Gauge |  | Variation within 0.015 mm
125 < MR £ 150
150 < MR £ 175
10 | Blue Matching |  | Blue and Gauge |  | More than 80% | 4 microns
200 < MR £ 225
225 < MR £ 250
11 | Parallelism |  | Height Gauge |  | 0.02 mm
5 | Thread Free Play |  | Visual |  | No perceptible free play (adjust threads if required)
6 | Constant Force Mechanism (if applicable) |  | Visual |  | Uniform Pressure
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## GRT Nick NG Masters
|  |  |  | Doc No : | ISW-QA-082
Musashi Auto Parts India Pvt Ltd |  |  |  | Issue No : | 1
|  |  |  | Rev No : | 0
|  |  |  | Date : | 2017-02-24
Calibration of GRT Nick NG Masters
Calibration Method: |  | ISW-QA-082
Calibration Freq.: |  | 1 month
1.0 Purpose: |  | This method describes the calibration of GRT Nick NG Masters.
2.0 Scope: |  | All GRT Nick NG Masters availble / required in Shop.
3.0 References : |  | Process Sheet.
4.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
5.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
6.0 Equipments Used:
Manual GRT, Master Gear & Dial Gauge 0.001 mm.
7.0 Calibration Process:
| 7.1 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 7.2 | Ensure the tracability of the master and equipment to National/International Standards.
| 7.3 | The test instrument should be throughly cleaned.
| 7.4 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for GRT NG Masters :
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
1 | Nick |  | Mount the Part on Manual GRT. Set the Dial Gauge on zero position. Complete one revolution of the part and observe the maximum value of Nick. Itch the Nick value thus observed on the Part. |  | For New Masters:
Nick Value (Itch Value)- (Component Nick Tolerance +0.010)±0.005
For Periodic Calibration:
±0.005 mm from the Itch Value
Identification display method:
White Paint Marker to identify Product Master.
Due Date Tepra Tape Sticker identification on GRT Nick NG Master as per method provided by MAP-ID.
Recorded Document:
Calibration Record of GRT Nick NG Master in GCS.
In case of Nick Master found NG during calibration Plese follow the rule "ISP-QA-010"
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Rolling Fixture
[5 empty] | Doc No : |  | ISW-QA-083
Musashi Auto Parts India Pvt Ltd |  |  |  |  | Issue No : | Issue No : | 1
[5 empty] | Rev No : | Rev No : | 0
[5 empty] | Date : | Date : | 2017-05-24
Calibration of Rolling Fixture
Calibration Method: |  | ISW-QA-083
Calibration Frequency: |  | 1 month & 3 month
1.0 Scope: |  | This method describes the calibration of Rolling Fixture.
2.0 References : |  | Gauge Drawing.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Slip Gauge Block, Plunger Dial and CMM.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 of 3
Permissible Deviations for Rolling Fixture :
Sr No | Test Characteristics |  | Measuring Instrument | Test Method |  |  |  | Acceptable Limit
1 | Center Height of Locators for Inspection Part Assembly |  | CMM | Check the center height of both the locators on CMM. |  |  |  | As per Drawing
2 | Diameter of locator for Inspection Part Assembly |  | CMM | Check the Diameter of both locators on CMM. |  |  |  | As per Drawing
3 | Alignment of Master Assembly Locator and Inspection Part Assembly Locator |  | Slip Gauge Block & Plunger Dial (0.001 mm) | Take a combination of gauge block equivalent to the gap between master assy. locator and inspection part assy. locator. Now place the gauge block combination between master assy locator and inspection part assy locator at right side and set the plunger dial at zero position. Next place the same gauge block combination between master assy locator and inspection part assy locator at left side and observe the diflection of plunger dail as the alignment error. |  |  |  | 0.020 mm
4 | Master Assembly Inspection |  | As per Rule ISW-QA-045 | Carry out the inspection of existing Master Assembly and New Master Assembly as per Rule ISW-QA-045 |  |  |  | As per Rule ISW-QA-045
5 | Pokayoke verification (if any) |  | As per Pokayoke Sheet
6 | LM Guide Movement |  | Hand feeling | Slide the LM Guide |  |  |  | Smooth and jerk free movement
7 | Bearing condition |  | Hand feeling | Rotate Master Assembly and feel the movement of Assembly |  |  |  | Smooth movement
8 | Master Assembly verification |  | OK Master Assembly | Mount the Master Assembly and rotate the assembly to verify the rolling and set the Plunger Dial at zero position. |  |  |  | Smooth rolling
page 2 of 3
Identification display method:
Colour Tepra Tape identification for 1 month frequency on Rolling Fixture as per method provided by MAP-ID.
Colour Paint Marker identification for 3 month frequency on Rolling Fixture as per method provided by MAP-ID.
Colour Coding
(1 month Frequency)
Colour Coding
(3 month Frequency)
Recorded Document:
Calibration Record of Rolling Fixture in GCS.
Tracibility:
Checked By
Prepared By
Approved By
Released By
NABL
EXTERNAL CALIBRATION AGENCY
SLIP GAUGE BLOCK, DIAL CALIBRATION TESTER
PLUNGER DIAL & DIAL TEST INDICATOR
ROLLING FIXTURE
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Go Pass Jig
|  |  |  | Doc No : |  | ISW-QA-084
Musashi Auto Parts India Pvt Ltd |  |  |  | Issue No : | Issue No : | 1
|  |  |  | Rev No : | Rev No : | 0
|  |  |  | Date : | Date :
Calibration of Go-Pass Jig 1 & Jig 2
Calibration Method: |  | ISW-QA-084
Calibration Frequency: |  | 1 month
1.0 Scope: |  | This method describes the calibration of Go-Pass Jig 1 & Jig 2
2.0 References : |  | Gauge Drawing.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Height Gauge, Dial Test Indicator and Three Pin Mic.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 of 3
Permissible Deviations for Rolling Fixture :
Sr No | Test Characteristics |  | Measuring Instrument | Test Method |  |  | Acceptable Limit
1 | Locator face to Gauge Plate Height of Jig-1 & Jig-2 |  | Dial Test Indicator & Height Gauge | Measure the height of gauge plate of Jig-1 & Jig-2 from the face of Locator. |  |  | As per Drawing
2 | Locator position verification |  | Setting block | Insert the setting block between locator side face and gauge plate. |  |  | Setting block should be in contact with locator and gauge plate
3 | Locator Diameter |  | Three Pin Mic | Measure the internal diameter of locator |  |  | As per Drawing
4 | Locator Height |  | Height Gauge | Measure the height of the locator. |  |  | As per Drawing
5 | Locator Chamfer |  | Visual | Check the availability of chamfer on locator internal diameter. |  |  | Chamfer should be available
6 | Pokayoke verification (if any) |  | As per Pokayoke Sheet
7 | LM Guide Movement |  | Hand feeling | Slide the LM Guide |  |  | Smooth and jerk free movement
8 | Visual Inspection |  | Visual | Check the condition of the Jig for dent, damage and burr. |  |  | No dent, damage and burr
9 | Master Assembly verification |  | Ok Master Assembly | Place the Ok master assembly on Jig-1 and slide the locator and pass through Gauge Plate. Similarly verify the Jig-2 with Ok master assembly. |  |  | Master Assembly should pass through the Gauge Plate.
page 2 of 3
Identification display method:
Colour Tepra Tape identification on Go-Pass Jig-1 & Jig-2 as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Go-Pass Jig-1 & Jig-2 in GCS.
Tracibility:
Checked By
Prepared By
Approved By
Released By
NABL
EXTERNAL CALIBRATION AGENCY
SETTING RING, DIAL CALIBRATION TESTER & CALIPER CHECKER
HEIGHT GAUGE, THREE PIN MIC & DIAL TEST INDICATOR
GO PASS JIG-1 & JIG-2
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Circlip Checking Gauge
|  |  |  | Doc No : |  | ISW-QA-085
Musashi Auto Parts India Pvt Ltd |  |  |  | Issue No : | Issue No : | 1
|  |  |  | Rev No : | Rev No : | 0
|  |  |  | Date : | Date :
Calibration of Circlip Checking Gauge
Calibration Method: |  | ISW-QA-085
Calibration Frequency: |  | 1 month
1.0 Scope: |  | This method describes the calibration of Circlip Checking Gauge
2.0 References : |  | Gauge Drawing.
3.0 Definations: |  | IS : Indian Standards
JIS:Japanese Industrial Standard
TI :Test Instruments
MR:Measuring Range
IMTE:Inspection Measuring and Test Equipment
4.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
5.0 Equipments Used:
Height Gauge, Dial Test Indicator and Three Pin Mic.
6.0 Calibration Process:
| 6.1 | The IMTE,Equipment and Master(if any) is ensured at Standard Temperature (20°±1°) for
an hour before calibration
| 6.2 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 6.3 | Ensure the tracability of the master and equipment to National/International Standards.
| 6.4 | The test instrument should be throughly cleaned.
page 1 of 3
Permissible Deviations for Rolling Fixture :
Sr No | Test Characteristics |  | Measuring Instrument | Test Method |  |  | Acceptable Limit
1 | Height of Go limit from Locator face (both side) |  | Dial Test Indicator & Height Gauge | Measure the height of Go limit from the face of Locator on both sides. |  |  | As per Drawing
2 | Height of NoGo limit from Locator face (both side) |  | Dial Test Indicator & Height Gauge | Measure the height of NoGo limit from the face of Locator on both sides. |  |  | As per Drawing
3 | Locator Diameter |  | Three Pin Mic | Measure the internal diameter of locator |  |  | As per Drawing
4 | Locator Height |  | Height Gauge | Measure the height of the locator. |  |  | As per Drawing
5 | LM Guide Movement |  | Hand feeling | Slide the LM Guide |  |  | Smooth and jerk free movement
6 | Visual Inspection |  | Visual | Check the condition of the Jig for dent, damage and burr. |  |  | No dent, damage and burr
7 | Circlip Master verification |  | Circlip Ok Master, Circlip Upside Master & Circlip Missing Master | Place all three master part one by one on locator and slide the locator and pass the master through Go and NoGo limit. |  |  | Ok Master: Pass through Go section but stuck in No Go
Circlip Upside Master: Will not enter the Go limit.
Circlip Missing Master: Pass through both Go and No Go section
page 2 of 3
Identification display method:
Colour Tepra Tape identification on Circlip Checking Gauge as per method provided by MAP-ID.
Colour Coding
Recorded Document:
Calibration Record of Circlip Checking Gauge in GCS.
Tracibility:
Checked By
Prepared By
Approved By
Released By
NABL
EXTERNAL CALIBRATION AGENCY
SETTING RING, DIAL CALIBRATION TESTER & CALIPER CHECKER
HEIGHT GAUGE, THREE PIN MIC & DIAL TEST INDICATOR
CIRCLIP CHECKING GAUGE
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By

## Spline Masters
|  |  |  | Doc No : | ISW-QA-086
Musashi Auto Parts India Pvt Ltd |  |  |  | Issue No : | 1
|  |  |  | Rev No : | 0
Date :
Calibration of Spline Masters
Calibration Method: |  | ISW-QA-086
Calibration Freq.: |  | 1 month
1.0 Purpose: |  | This method describes the calibration of Spline Masters.
2.0 Scope: |  | This document describes the calibration procedure for all Spline Product Masters used to set
the limit of Spline Arbor and all Spline Product Masters used to verify Citizen Gauge in 4w.
3.0 References : |  | Process Sheet, Final drawing of component.
4.0 Definations: |  | IS : Indian Standards
JIS: Japanese Industrial Standard
TI : Test Instruments
MR: Measuring Range
IMTE: Inspection Measuring and Test Equipment
5.0 General Requirements:
Environment:
| 1) | Temperature: 20°C
| 2) | Humidity : No Excessive Humidity
| 3) | Air Quality : N/A
Stabilization: |  | Stablize equipment and standards at ambient temperature for two hours (Min)
Preliminary Operations:
| 1) | Clean TI.
| 2) | Verify TI for damage such as nicks or burrs.
Standards and Calibrating Equipment:
The standards and equipment used must have a valid calibration certificate.
6.0 Equipments Used:
Citizen Gauge & Dial Gauge 0.001 mm.
7.0 Calibration Process:
| 7.1 | Ensure the L.C of equipment used for calibration is at least 1/10th of the TI.
| 7.2 | Ensure the tracability of the master and equipment to National/International Standards.
| 7.3 | The test instrument should be throughly cleaned.
| 7.4 | To obtain the repeatability 3 to 5 observations are taken for the particular parameter.
page 1 of 2
Permissible Deviations for Spline Masters :
1) Spline Masters for Spline Arbor
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
1 | Spline Major Dia |  | Mount the part on citizen gauge and inspect the major dia of the part on every teeth and take the average value or minimum value as per method defined in process sheet. Next itch the Max or Min on the part as per the observation of the part. |  | For Min Major Dia Master:
Master Value = Lower value of Spec. of Major Dia. + 0.005 mm
For Max Major Dia Master:
Master Value = Higher value of Spec. of Major Dia. - 0.005 mm
2) Spline Masters for Citizen Gauge Verification
Sr No | Test Characteristics |  | Test Method |  | Acceptable Limits of Error
1 | Spline Major Dia |  | Mount the part on citizen gauge and inspect the major dia at any teeth at the height specified in Final Drawing of component and observe the value of Major Dia. Mark the teeth on which the Inspection is carried out and itch the value observed on the part. |  | For Making New Master:
Itch Value = Major Dia spec. specified in Final Drawing
For Periodic Master:
Itch Value ± 0.002 mm
Identification display method:
White Paint Marker to identify Product Master.
Due Date Tepra Tape Sticker identification on Spline Master as per method provided by MAP-ID.
Recorded Document:
Calibration Record of Spline Masters in GCS.
Checked By
Prepared By
Approved By
Released By
Checked By
Prepared By
Approved By
Released By